## Slide 1
Advanced Flow Cytometry
J. Paul Robinson & John Nolan

## Slide 2: 3-Spectral Flow Cytometry
John P. Nolan
Scintillon Institute
San Diego, CA
jnolan@scintillon.org
Advanced Short Course

## Slide 3: Learning Objectives
Principles of spectral flow cytometry
Instrument approaches
Applications of spectral flow cytometry

## Slide 4: Flow Cytometer Schematic
Sample Inlet
Sheath Inlet
Beam Shaping Optics
Sample
Stream
Sheath
Stream
Filters
Dichroic mirrors
Detectors
Data Acquisition System
Collection
Optics

## Slide 5: Polychromatic Flow Cytometry
Perfetto, 2004
Notes: Instrumentation for multiparameter measurements
Violet, blue, green, and red lasers
Efficient light collection and spectral separation
Cross talk occurs, dealt with through compensation
System is maxed out from and instrument standpoint, but what about reagents?

## Slide 6: Spectral FC: The first report…

## Slide 7: Spectral FC: Grating and CCD

## Slide 8: Diffraction grating and multianode PMT

## Slide 10: Sony Spectral Analyzer

## Slide 11: Scintillon Spectral Flow Cytometer
FALS Detector
Flow Cell
FALS Collection
SSC/ Fluorescence/
Raman
Collection
Spectrograph/ CCD
Changeable gratings
Adjustable slits
On-chip pixel binning
PMT/Filter Assembly
Laser beams
405 nm
488 nm
660 nm
Top View
Front View
Performance - spectral
Spectral axis: 1600 pixels
Spectral resolution: 0.3 nm (max)
Speed: ~200 events/sec (max)
Next gen:~2000 events/sec (max)
Performance - conventional
Sensitivity: 27 MESF PE
Separation limit: 81 MESF PE

## Slide 12: Photodetectors
Multianode PMT
Charge coupled device (CCD)
Fast
32 channels
Low QE
High QE
1600 channels
Slow

## Slide 13: Spectral FC Data Analysis

## Slide 14: Instrument Performance
Multifluorophore, multi-intensity reference particles – Spherotech UltraRainbow
Single fluorophore, multi-intensity calibration particles – BD Quantibrite PE
~60 mW 488 nm, 25 usec pulse width
Nolan et al (2013) Cytometry

## Slide 15: PBMC Immunophenotyping
Antibody Capture Beads
Single Color Reference Spectra
Nolan et al (2013) Cytometry

## Slide 16: PBMC Immunophenotyping
Cells
CD3+
Raw Spectra
Reference Spectra
(capture beads)
Unmixed Signals
Nolan et al (2013) Cytometry
Spectral
Unmixing

## Slide 17: PBMC Immunophenotyping
Cells
CD3+
CD14+
CD3+
Cells
CD3+, CD4+
CD3+, CD8+
Nolan et al (2013) Cytometry

## Slide 18: PBMC Immunophenotyping
Cells
Cells
Cells
CD3+
CD3+
CD14+
CD3+
Nolan et al (2013) Cytometry

## Slide 19: Unmixing autofluorescence
Sony Biotechnology

## Slide 20: Unmixing autofluorescence
Sony Biotechnology

## Slide 21: Summary
Spectral flow cytometry measures all of the emission from cells
Spectral unmixing replaces compensation
Performance is equivalent or better than conventional flow cytometry
Autofluorescence becomes an additional parameter rather than a source of background