## Slide 1
Advanced Flow Cytometry
J. Paul Robinson & John Nolan

## Slide 2: 4- Mass Cytometry – A new Technology Impacting Flow Cytometry
J. Paul Robinson
The SVM Professor of Cytomics
Professor of Biomedical Engineering
Director, Purdue University Cytometry Laboratories
Advanced Short Course
I wish to thank Dr. Hervé Luche from the
Centre d’Immunophenomique Marseille, France
For providing most of the slides on mass cytometry

## Slide 3
Key Links
cyto.purdue.edu
isac-net.org
scintillon.org/nolan_lab
Purdue List Serve &
Robinson Lab
ISAC
Nolan Lab
John P. Nolan - jnolan@scintillon.org
J. Paul Robinson – wombat@purdue.edu

## Slide 4: What is the issue with Fluorescence based Flow Cytometry?
Wave Length (nm)
Emission spectra
There are a limited number of bands that we can handle using fluorescence
What are the alternatives?

## Slide 5: The immuno-technology behind mass cytometry…
Isotopic Mass (Da)
4-5 polymer per antibody
16-22 chelating unit per polymer
100-80 atoms coupled per antibody
MaxPAR Antibodies
Syrigen dye – Antibody
Wave Length (nm)
Emission spectra

## Slide 6: ….allied to the analytical power of ICP-MS
135 channels
45 mAb Tags
6 Dyes
Helios
www.editions-petiteelisabeth.fr/tableau_periodique_des_elements_2.php

## Slide 7: Combined immuno-phenotyping and functional workflow

## Slide 8: Mass cytometry principle

## Slide 9: How does the raw data look like ?
Newel, Immunity, 2012

## Slide 10: How to immuno-phenotype ?
Variations
Operator
Reagents
Instrument
Data Analysis
NEUTRAL
I
ZE

## Slide 11
Experiment: 13-SE19-TEST-MMK
1µg of CD45.2-141Nd
CD45.2-141Nd-2013JL26
61 atoms/mAb
Intensity signal : 63
1µg of CD45.2-156Gd
CD45.2-156Gd-2013SE03
61 atoms/mAb
Intensity signal : 342
1µg of CD45.2-176Yb
CD45.2-176Yb-2013JL05
47 atoms/mAb
Intensity signal : 207
The difference in sensitivity is really minimal
Same trend observed fitting with intrument sensitivity curves but this has no impact on the frequencies observed. Even in the worst combination, we can accurately measure the cell frequency of each genotype

## Slide 12: Implication for panel design
All metals are not equals in terms of sensitivity
As for other cytometry method, panel design is required
Tricot et al., Cytometry A (2015)

## Slide 13: Systematic validation method for new antibodies/metal conjugates
    Staining intensity
The stain index in mass cytometry is less than conventional fluorescent dyes

## Slide 14: Main types of contamination in mass cytometry
146Nd
145Nd
Impurities (-1)
Abondance (+1)
Oxydes (+16)
Signal Cross-talks are due to:
Isotopic impurities of metals (+ and/or – several Da)
Abundance sensitivity (+ and/or – 1 Da)
Oxidation (+16 Da) of reporter ions during the acquisition
Takahashi, Cytometry Part. A, Cytometry A. 2016 Sep 15

## Slide 15: Cold case: impurity issue
Unstimulated
Stimulated
LatY136F CD4 Tc

## Slide 16
Sample preparation is key: Effect of inappropriate initial fixation
for intranuclear staining

## Slide 17
Molecules below 100 molecules/cell cannot be identified due to ion detection efficiency (1 in 10.000 ion)
30% (CyTOF1) to 60% (Helios) of the cells injected are actually acquired by the system
All metals are not equals (2-3 X difference between chanels). This difference in sensitivity is not in the same range than flow (40-50 X)
Dynamic range of measurement is smaller than by fluoresence flow cytometry
No spectral overlaps but cross-talks and impurities that can spoil your measured signals
Carefully design marker/metal pair
Carefully choose conjugates combinations for a meaningful analysis
Hudge responsability in the field of mass cytometry to seriously consider this step :
  unprecedented combinations of cellular marker investigated: QC ?
  Unravel heterogeneity that is real not artefactual

## Slide 18: Strategy for careful panel design
Basal
Stimulated
Validate all you protocols by conventional flow
Establish ranking of marker you want to adress in one chanel
Use your mass response curve as a normaliser to predict signal intensity in a given chanel
Use Panel designer tool to minimize crosstalks

## Slide 19: How to apprehend the multi-parametric headache
1983
Parameters:0004
Plots:0006
1989
Parameters:0008
Plots:0028
2000-2010
Parameters:0014
Plots:0091
2011
Parameters:0032
Plots:0496
2016
Parameters: 40
Plots: >600

## Slide 20: Unsupervised analysis methods to handle this amount of data
Levine, Cell, 2015
Cheng, JI, 2015
OneSense
ViSNE
FlowSOM
Van Gassen, Cytometry, 2015

## Slide 21: Concept of analysis pipelines
Diggins, Methods, 2015

## Slide 22: Mass/flow correlation for CD8 Tc sub-populations
Chart: 
Chart: 
| 98.8 | 58.3 | 31 | 2.1 | 97.8 | 1.2 | 65.3 | 60.1 | 22.6 | 43.4 | 33.700000000000003 | 50.3 | 4.8 | 7 | 20.7 | 43.2 | 7 | 24.8 | 12.4 | 53.5 | 15.6 | 6.9 | 66.900175131348533 | 68.014059753954299 | 31.985940246045679 | 33.099824868651467
 | 98.44 | 54 | 27.28 | 2.61 | 97.1 | 1.24 | 64.46 | 57.44 | 25.18 | 44.3 | 32.14 | 57.71 | 6.56 | 7.12 | 23.87 | 46.79 | 6.97 | 14.29 | 13.09 | 55.45 | 12.91 | 3.27 | 69.5013 | 69.5013 | 30.4987 | 30.4987
Mass
Flow
CD8+ L-T

## Slide 23: How to setup a CyTOF experiment ?
Configure your instrument (laser power, BP/LP/ND filters)
Characterize your instrument (CSTs and QCs)
Design your panel accordingly
Optimize settings for your assay (2,5 SDEN, Lin Max)
Run appropriate controls (Comp Beads, FMOs for continuums)
QC your data (Dead cells, Doublets, Clogs, Mis-annotations) !
Use adapted analysis tools
Configure your instrument (QC, Mass Range, Dual Calibration)
Characterize your instrument (Sensitivity across Mass range)
Optimize settings for your assay (Nebular Flow for Oxydation)
Design your panel accordingly
Run appropriate controls (Eu/eQ/Comp Beads, Cells + DNA)
QC your data (Dead cells, Ion Cloud Fusion, Doublets)
Use adapted analysis tools (Cytobank + SPADE, VisNE)

## Slide 24: Available tools to help in mass cytometry
Panel Kits
Panel Designer tool
Normalisation Beads

## Slide 25: A rapidly evolving suite of instruments
2011
CyTOF1
2013
CyTOF2
2015
Helios

## Slide 26: Proliferation assay made simple: IdU incorporation and detection on CyTOF
hCD45+ JURKAT cells
Incubation time
1%
35%
46%
55.7%
IdU detection is easier than BrdU assay by conventional flow

## Slide 27
Splenocytes
Activated Tc
« Proliferation » gate
« Singlets » gate
Splenocytes
Proliferating cells
Classical gating method is not appropriate
65
Loss of proliferating cells

## Slide 28: Sample multiplexing via Barcoding
Cells are encoded with combinations of various dyes
Each combination is unique and refers to a given well
All wells can later be combined for simultaneous staining and acquisition
Barcode 1
Barcode 2
Barcode 3
CD45 (30F11)
CD45.2 (104)
Thymus/Spleen/LN

## Slide 29: Sample Barcoding for multiplexed analysis
Minimize experimental variations
Decrease experimental artefacts
Decrease cell loss
Reduce time at washing steps
Allow simultaneous analysis
Improved data comparison

## Slide 30: Acquisition and standardized analysis
CD8
Ly6G/C
MHCII
CD4
All frequencies are determined on similar VISNE plot for all mutants
Same panel used on all samples at once
Antibody staining done one the same day for all samples
Same analysis strategy can be used

## Slide 31: How does SPADE work ? (Spanning-tree Progression Analysis of Density-normalized Events)
Qiu, Nat. Biotech, 2011, “Extracting a cellular hierarchy from high-dimensional cytometry data with SPADE”
SPADE (Spanning-tree Progression Analysis of Density-normalized Events) is a way to automatically identify populations in multidimensional flow cytometry data files.

## Slide 32: Unsupervised analysis and visualization using SPADE trees
CTRL
CTRL
MUT
CD3
CD4
MUT

## Slide 33: Cytokine and Transcription factor detection in dLN and tumor infiltrating leukocytes
CD4 Tc
CD8 Tc
IL2
IL4
IL17A
IFNg
IL2
TNFa
CD4+ TILs

## Slide 34: PlateAnalyzer from Purdue can also analyze mass cytometry data
http://www.cyto.purdue.edu/Purdue_software
14 Phospho markers
14 lymphoid markers
12 Act’ Molecules/8 conc

## Slide 35: The Bottom line – There are alternatives to traditional flow cytometry
Mass Cytometry is
Mass Cytometry provides similar results to flow cytometry
It uses the same antibodies, but they are conjugated with heavy metals not fluorescent molecules
The data analysis is more complex than flow cytometry as there are usually more parameters
Flow cytometry has demonstrated an increased number of parameters probably because of mass cytometry
There are advantages of both technologies as there are disadvantages