## Slide 1
Advanced Flow Cytometry
J. Paul Robinson & John Nolan

## Slide 2: 5-Exosomes and Nanoparticles
John P. Nolan
Scintillon Institute
San Diego, CA
jnolan@scintillon.org
Advanced Short Course

## Slide 3: Learning Objectives
Biological nanoparticles
  Exosomes and microvesicles
Differences between cells and vesicles
Detection approaches
Calibration and data reporting

## Slide 4: All cells release small vesicles

## Slide 5: In biofluids, vesicles come from many different cells

## Slide 6: EV Analysis Opportunities and Challenges
Biomarkers
Therapeutic targets
Therapies
EVs are heterogeneous
  Shed from the surface or secreted
  Come from many different cell types
  Require multiparameter analysis
EVs are small and dim
  They don’t scatter much light
  They don’t have many antigens
  Conventional FC performs poorly

## Slide 7
Cell
10 um diameter
CD16: 10,000 copies (50/um2)
CD11a: 55,000 copies (274/um2)
CD14: 110,000 copies (547/um2)
1 um
100 um
Cell-derived EVs
80 - 400 um diameter

## Slide 8: EV Analysis Methods
Bulk Analysis
  Dynamic light scattering
    Size distribution
  Immunoassay
    Total antigen
  Genomics
    miRNA
  Proteomics
    Mass spec
Single Particle Analysis
  Electron microscopy, cryo EM, AFM
    Size, immuno-gold
  Nanoparticle tracking analysis
    Size, intensity
  Resistive pulse spectroscopy
    Size, zeta potential
  Flow cytometry
    Size (?), multiple antigens
EV Isolation – Blood, urine, saliva, CSF, culture media
  Differential/gradient/ultra centrifugation, filtration, precipitation

## Slide 9: Microscopies
Cryo EM (+immunogold)
Freeze fracture EM
AFM

## Slide 11: Nanoparticle Tracking Analysis (NTA)
Individual vesicles are detected via laser light scatter
Brownian motion tracked
Size estimated from diffusion constant
Nanosight.com

## Slide 12: Resistive Pulse Spectroscopy (RPS)
Coulter principle using nanopores
Particles in an electrolyte solution impede current when they enter pore
Magnitude of impedance is proportional to size
www.izon.com

## Slide 13: How big are EVs?
~70-200 nm diameter as measured by many different methods
Many exceptions:
oncosomes
red cell ghosts/fragments
tubules
are minority
Varga et al 2014 J Extracell Ves

## Slide 14: Why flow cytometry?
EVs are heterogeneous
  Ectosomes, exosomes, others
Biological fluids contain EVs from different cells
Multiparameter analysis will distinguish these
Enumeration (EVs/mL)
Marker quantification (molecules/EV)

## Slide 15: Bead-based Flow Cytometry
Microspheres functionalized with anti-exosome capture antibodies
  CD9, CD63, MHC II
Captured EVs labeled with fluorescent anti-exosome antibody
Bead fluorescence measured by flow cytometry
Kinker et al (2014) Frontiers in Immunology

## Slide 16: Single EV Analysis: “Calibrated” Size Discrimination
“ Microparticles are defined as small (0.1-1 um) diameter anucleoid phospholipid vesicles released from different cells …[ref]“

## Slide 17: Facts of Light Scatter
Light scatter does not measure size (or granularity)
Light scatter can, under certain circumstances, be proportional to size (or granularity)
Light scatter intensity is a complex function of wavelength, particle size and shape, angle of collection, and refractive index
Light scatter is well described mathematically
www.philiplaven.com/mieplot.htm

## Slide 18: Angle of Collection and Size Calculated for 488 nm excitation
www.philiplaven.com/mieplot.htm
Above λ, light scatter is highly angle dependent
LS is much stronger at small (forward angles)
background is also higher (obscuration)
Below λ, particles behave as points, and little angular dependence is observed
- Lower background makes 90⁰ better for measurement of small particles

## Slide 19: Particle Diameter Calculated at 488 nm excitation and 80-100⁰ collection
Above λ, monotonic increase in scatter with increasing radius (but be careful with FALS)
Below λ, scatter decreases with an r6 dependence (small particles get very dim, very fast)
www.philiplaven.com/mieplot.htm

## Slide 20: Refractive Index and Size Calculated for 488 nm excitation and 80-100⁰ collection
www.philiplaven.com/mieplot.htm
Light scatter from a lipid vesicle is predicted to be >100 fold lower compared to a polystyrene bead

## Slide 21: Reference Particles
Polymer beads have a long history of use for standardization and calibration of FC measurements, however caution is warranted in using them for light scatter measurements
Robert et al 2009 Journal of Thrombosis and Haemostasis

## Slide 22: ISTH Standardization Effort

## Slide 23: How big are EVs?
van der Pol et al 2010 J Thromb Haem
Comparison of various methods to
Small Angle X-ray Scattering (SAXS)

## Slide 24: FC of Individual EVs: Pitfalls
Lack of sensitivity
Artifacts!!!
Irreproducibility
Difficulty of standardization
Lack of appropriate calibration

## Slide 25: Coincidence?
probability of having n particles within the probe volume simultaneously
where h is the height of the laser beam, R is the event rate, and v is the linear flow velocity
Notes: The probability of coincident events can be calculated using the following equation: where is the probability of having n particles within the probe volume simultaneously, and Г is defined by the following relation: , where h is the height of the laser beam, R is the event rate, and v is the linear flow velocity. Coincidence calculations were limited calculating the probability of a doublet (2 coincident particles).

## Slide 26
Cytometry Part A 16 FEB 2015 DOI: 10.1002/cyto.a.22644 http://onlinelibrary.wiley.com/doi/10.1002/cyto.a.22644/full#cytoa22644-fig-0001
High concentrations of small dim particles can result in coincidence artifacts (aka “swarm”)
Notes: Coincident detection of 100 nm fluorescent beads at high bead concentrations. Serial dilutions from a highly concentrated 100 nm bead suspension were prepared and analyzed using a 140 μm nozzle and 5 psi sheath fluid pressure. (A) Analysis of event rate, (B) light scattering (rw‐FSC) and (C) fluorescence (FL1). (D) Scatter plots (rw‐FSC vs. FL1) of bead samples indicated by arrows in A–C. Scatter plots display 20,000 recorded events. Below are images of the digital oscilloscope captured during the analysis of low bead concentrations (left), during coincident particle detection (middle), and during clear swarm (right). Yellow and blue arrows indicate the level of the baseline signal. Shown is a representative experiment out of three experiments. [Color figure can be viewed in the online issue, which is available at wileyonlinelibrary.com.]
© IF THIS IMAGE HAS BEEN PROVIDED BY OR IS OWNED BY A THIRD PARTY, AS INDICATED IN THE CAPTION LINE, THEN FURTHER PERMISSION MAY BE NEEDED BEFORE ANY FURTHER USE. PLEASE CONTACT WILEY'S PERMISSIONS DEPARTMENT ON PERMISSIONS@WILEY.COM OR USE THE RIGHTSLINK SERVICE BY CLICKING ON THE 'REQUEST PERMISSION' LINK ACCOMPANYING THIS ARTICLE. WILEY OR AUTHOR OWNED IMAGES MAY BE USED FOR NON-COMMERCIAL PURPOSES, SUBJECT TO PROPER CITATION OF THE ARTICLE, AUTHOR, AND PUBLISHER.

## Slide 27: Coincidence detection: Dilution

## Slide 28: Limitations of Conventional FC
Light scatter as a trigger channel
  Depends on size, shape, λ, collection angle, refractive index
  Well described by Mie theory
  Vesicles scatter 10-100x <beads
Coincidence
  Depends on [EV], probe volume
  Frequency is readily calculated
  Can be identified/eliminated by dilution
Fluorescence calibration
  Required for data/methods sharing
  Well-established protocols, reagents
  No one does it
PS Beads
Vesicles

## Slide 29
Fluorescence triggering based on surface marker staining

## Slide 31: Vesicle Flow Cytometry
Stained sample
Cell-depleted biofluid
Spin@ 2,500xg (2X) to remove cells, large debris
Stain with membrane probe, antibody
High sensitivity flow cytometry

## Slide 32: Vesicle Flow Cytometry
Lipid Probe:
EV Surface area
calibrated using liposomes
Marker Probe:
Marker abundance estimated through fluorescence calibration
Bivariate plot:
correlated data
All particles introduced into FC
Only membranous particles (EVs) are detected and measured
Instrument calibration enable estimation of surface marker abundance
1980 AnnV
molecules

## Slide 33: Summary
Extracellular vesicles and other biological nanoparticles are potentially important
Nanoparticles are not cells,
Accurate measurement of small, dim particles requires different approaches
Calibration and controls are essential

## Slide 34: Resources
Current Protocols in Cytometry
Unit 1.3 Standardization, Calibration, and Control in Flow Cytometry
Unit 1.4 Establishing and Maintaining System Linearity
Unit 1.20 Characterization of Flow Cytometer Instrument Sensitivity
Unit 6.4 Enumeration of CD34+ Hematopoietic Stem and Progenitor Cells
Unit 6.8 Enumeration of Absolute Cell Counts Using Immunophenotypic Techniques
Unit 6.24 Flow Rate Calibration for Absolute Cell Counting Rationale and Design
Unit 6.26 Calibration of Flow Cytometry for Quantitative Quantum Dot Measurements
Unit 13.2 Microsphere Surface Protein Determination Using Flow Cytometry
Unit 13.14 Flow Cytometry of Extracellular Vesicles: Potential, Pitfalls, and Prospects
CYTO 2013 Scientific Tutorial:
Cytometer Performance Characterization and Standardization Robert Hoffman
CYTO U Webinar:
Predicting the best resolution and sensitivity in panel development and reducing inter-instrument variability in flow cytometry. Steve Perfetto and Jim Wood

## Slide 35: More Resources

## Slide 36: Protocols and More Info
Chapter 1 Flow Cytometry Instrumentation
Chapter 2 Image Cytometry Instrumentation
Chapter 3 Safety Procedures and Quality Control
Chapter 4 Molecular and Cellular Probes
Chapter 5 Specimen Handling, Storage, and Preparation
Chapter 6 Phenotypic Analysis
Chapter 7 Nucleic Acid Analysis
Chapter 8 Molecular Cytogenetics
Chapter 9 Studies of Cell Function
Chapter 10 Data Processing and Analysis
Chapter 11 Microbiological Applications
Chapter 12 Cellular and Molecular Imaging
Chapter 13 Multiplexed and Microparticle‐Based Analyses
  Introduction
  UNIT 13.1 Multiplexed Microsphere-Based Flow Cytometric Immunoassays
  UNIT 13.2 Microsphere Surface Protein Determination Using Flow Cytometry
  UNIT 13.3 Use of Microsphere-Supported Phospholipid Membranes for Analysis of Protein-Lipid Interactions
  UNIT 13.4 Multiplexed SNP Genotyping Using Primer Single-Base Extension (SBE) and Microsphere Arrays
  UNIT 13.5 BeadCons: Detection of Nucleic Acid Sequences by Flow Cytometry
  UNIT 13.6 Characterization of Nuclear Receptor Ligands by Multiplexed Peptide Interactions
  UNIT 13.7 Detection of Gene Fusions in Acute Leukemia Using Bead Microarrays
  UNIT 13.8 Reagents and Instruments for Multiplexed Analysis Using Microparticles
  UNIT 13.9 Multiplexed Detection of Fungal Nucleic Acid Signatures
  UNIT 13.10 Multiplexed Analysis of Peptide Antigen-Specific Antibodies
  UNIT 13.11 Use of Flow Cytometric Methods to Quantify Protein-Protein Interactions
  UNIT 13.12 Microsphere-Based Flow Cytometry Protease Assays for Protease Activity Detection & High-Throughput Screening
  UNIT 13.13 Application of the PrimRglo Assay Chemistry to Multiplexed Bead Assays
  UNIT 13.14 Flow Cytometry of Extracellular Vesicles: Potential, Pitfalls, and Prospects
  UNIT 13.15 Optimized MOL-PCR for Characterization of Microbial Pathogens