## Slide 1: BMS 524 - “Introduction to Confocal Microscopy and Image Analysis”
Lecture 5: Fluorescence
Department of Basic Medical Sciences,
College of Veterinary Medicine
and
Weldon School of Biomedical Engineering
College of Engineering
Purdue University
J. Paul Robinson, Ph.D.
Distinguished Professor of Cytometry
& Professor of Biomedical Engineering
Director, Purdue University Cytometry Laboratories, Purdue University
This lecture was last updated in January, 2023
Find other PUCL Educational Materials at http://www.cyto.purdue.edu/class
These slides are intended for use in a lecture series. Copies of the slides are distributed and students encouraged to take their notes on these graphics. All material copyright J.Paul Robinson unless otherwise stated. No reproduction of this material is permitted without the written permission of J. Paul Robinson. Except that our materials may be used in not-for-profit educational institutions with appropriate acknowledgement. It is illegal to upload this lecture to CourseHero or any other site.

## Slide 2: Overview
Fluorescence
The fluorescent microscope
Types of fluorescent probes
Problems with fluorochromes
General applications

## Slide 3: Learning Objectives – Lecture 5
At the conclusion of this lecture, you should:
Understand the nature of fluorescence
The restrictions under which fluorescence occurs
Nature of fluorescence probes
Spectra of different probes
Resonance Energy Transfer and what it is
Features of fluorescence

## Slide 4: Excitation Sources
Excitation Sources
Lamps
    Xenon
    Xenon/Mercury
Lasers
    Mostly solid-state lasers
    Argon Ion (Ar) Krypton (Kr) (almost never used)
    Blue 488 nm
    Violet 405nm, 380 nm
    Green 561 nm
    Helium-Neon (He-Ne)
    Helium-Cadmium (He-Cd) (rarely used)
    Krypton-Argon (Kr-Ar) (no longer used)
Laser Diodes
375nm up to NIR
LEDs see next side
2004 sales of approximately 733 million diode laser; 131,000 of other types of lasers

## Slide 5: LEDs in Microscopy
Photon-emitting diode p-n junctions are typically based on a mixture of Group III and Group V elements, such as gallium, arsenic, phosphorous, indium, and aluminum. The relatively recent addition of silicon carbide and gallium nitride to this semiconductor palette has yielded blue-emitting diodes, which can be combined with other colors or secondary phosphors to produce LEDs that emit white light. The fundamental key to manipulating the properties of LEDs is the electronic nature of the p-n junction between two different semiconductor materials. When dissimilar doped semiconductors are fused, the flow of current into the junction and the wavelength characteristics of the emitted light are determined by the electronic character of each material.
https://www.zeiss.com/microscopy/us/solutions/reference/all-tutorials/light-sources/led-operation.html
https://www.electronicshub.org/led-light-emitting-diode/

## Slide 6: Higher Capacity by Smaller Spot and Thinner Cover
DVD
BD
CD
Wavelength: 650 nm
NA : 0.60
Capacity: 4.7GB
D = 0.88um
Wavelength: 405 nm
NA : 0.85
Capacity: 25GB
D = 0.39um
Wavelength: 780 nm
NA : 0.45
Capacity: 0.78 GB
Spot Size D = 1.42um
Cover Thickness
1.2mm
Cover Thickness
0.6mm
Cover Thickness
0.1mm
Pit Mastering 442nm He-Cd
406nm Kr
413nm Ar
405nm PTM
E-beam
257nm Ar
350nm Ar/Kr
780nm
650nm
405nm

## Slide 7: Fluorescence

## Slide 8: Fluorescence
Review:
Chromophores are components of molecules which absorb light
e.g. from protein most fluorescence results from the indole ring of tryptophan residue
They are generally aromatic rings

## Slide 9: Fluorescence
The wavelength of absorption is related to the size of the chromophores
Smaller chromophores, higher energy (shorter wavelength)

## Slide 10: Fluorescence
Stokes Shift
  is the energy difference between the lowest energy peak of absorbance and the highest energy of emission
495 nm
518 nm
Stokes Shift is 25 nm
Fluorescein
molecule
Fluorescence Intensity
Wavelength

## Slide 11: Fluorescence
The longer the wavelength the lower the energy
The shorter the wavelength the higher the energy
  e.g. UV light from sun causes the sunburn
  not the red visible light
Intensity
related to the probability of the event
Wavelength
  the energy of the light absorbed or emitted

## Slide 12: First Synthetic Dyes
https://www.ncbi.nlm.nih.gov/pmc/articles/PMC2924670/

## Slide 13: Light Sources - Lasers
Argon Ar 353-361, 454, 488, 514nm
Violet Diode 380-405 nm
Krypton-Ar Kr-Ar 488, 568, 647nm
Helium-Neon He-Ne 543 nm, 633nm
He-Cadmium He-Cd 325 or/and 441nm
Diode – (CD) 780nm
Diode – (DVD) 650nm
Diode – (Blu-Ray) 405nm
Solid-State Many wavelengths
Laser Abbrev. Excitation Lines
(He-Cd light difficult to get 325 nm band through some optical systems – need quartz)

## Slide 14: Arc Lamp Excitation Spectra
Irradiance at 0.5 m (mW m-2 nm-1)




   
Xe Lamp
Hg Lamp
488
632
405
350

## Slide 15
Xenon Lamp
Mercury Lamp
Xenon/Mercury Lamp
This is why Xenon is used

## Slide 16: Excitation - Emission Peaks
Fluorophore EXpeak EMpeak
% Max Excitation at
488 568 647 nm
FITC 496 518 87 0 0
Bodipy 503 511 58 1 1
Tetra-M-Rho 554 576 10 61 0
L-Rhodamine 572 590 5 92 0
Texas Red 592 610 3 45 1
CY5 649 666 1 11 98
Note: You will not be able to see CY5 fluorescence
under the regular fluorescent microscope because
the wavelength is too high.
Material Source:
Pawley: Handbook of Confocal Microscopy
Note: the 3 laser combination is no longer used

## Slide 17: Parameters
Extinction Coefficient
   (extinction coefficient) refers to a single wavelength (usually the absorption maximum)
Quantum Yield
  Qf is a measure of the integrated photon emission over the fluorophore spectral band
At sub-saturation excitation rates, fluorescence intensity is proportional to the product of  and Qf
Number of emitted photons
Number of absorbed photons
=
Lifetime 1 –10x10-9secs (1-10 ns)

## Slide 18: Fluorescence Lifetimes
https://www.ncbi.nlm.nih.gov/pmc/articles/PMC2924670/

## Slide 19: Fluorescence Lifetimes
http://www.iss.com/resources/reference/data_tables/LifetimeDataFluorophores.html

## Slide 20: Relative absorbance of phycobiliproteins
Protein | 488nm % absorbance | 568nm% absorbance | 633nm % absorbance
B-phycoerytherin | 33 | 97 | 0
R-phycoerytherin | 63 | 92 | 0
allophycocyanin | 0.5 | 20 | 56
Data from Molecular Probes Website
Phycobiliproteins are stable and highly soluble proteins derived from cyanobacteria and eukaryotic algae with quantum yields up to 0.98 and molar extinction coefficients of up to 2.4 × 106

## Slide 21: Excitation Saturation
The rate of emission is dependent upon the time the molecule remains within the excitation state (the excited state lifetime f)
Optical saturation occurs when the rate of excitation exceeds the reciprocal of f
In a scanned image of 512 x 768 pixels (400,000 pixels) if scanned in 1 second requires a dwell time per pixel of 2 x 10-6 sec.
Molecules that remain in the excitation beam for extended periods have higher probability of interstate crossings and thus phosphorescence
Usually, increasing dye concentration can be the most effective means of increasing signal when energy is not the limiting factor (ie laser based confocal systems)
Material Source:
Pawley: Handbook of Confocal Microscopy

## Slide 22: How many Photons?
Consider 1 mW of power at 488 nm focused to a Gaussian spot whose radius at 1/e2 intensity is 0.25m via a 1.25 NA objective
The peak intensity at the center will be 10-3W [.(0.25 x 10-4 cm)2]= 5.1 x 105 W/cm2 or 1.25 x 1024 photons/(cm2 sec-1)
At this power, FITC would have 63% of its molecules in an excited state and 37% in ground state at any one time
C21H11NO5S
Material Source:
Pawley: Handbook of Confocal Microscopy

## Slide 23: Photobleaching
Defined as the irreversible destruction of an excited fluorophore (discussed in later lecture)
Methods for countering photobleaching
  Scan for shorter times
  Use high magnification, high NA objective
  Use wide emission filters
  Reduce excitation intensity
  Use “antifade” reagents (not compatible with viable cells)

## Slide 24: Quenching
Not a chemical process
Dynamic quenching =- Collisional process usually controlled by mutual diffusion
Typical quenchers – oxygen
Aliphatic and aromatic amines* ( NO2, CHCl3)
Static Quenching
Formation of ground state complex between the fluorophores and quencher with a non-fluorescent complex (temperature dependent – if you have higher quencher ground state complex is less likely and therefore less quenching)
*primary amines - only one alkyl group

## Slide 25: Antifade Agents
Many quenchers act by reducing oxygen concentration to prevent formation of singlet oxygen
Satisfactory for fixed samples but not live cells!
Antioxidents such as propyl gallate, hydroquinone, p-phenylenediamine are used
Reduce O2 concentration or use singlet oxygen quenchers such as carotenoids (50 mM crocetin or etretinate in cell cultures); ascorbate, imidazole, histidine, cysteamine, reduced glutathione, uric acid, trolox (vitamin E analogue)

## Slide 26: Photobleaching example
FITC - at 4.4 x 1023 photons cm-2 sec-1 FITC bleaches with a quantum efficiency Qb of 3 x 10-5
Therefore, FITC would be bleaching with a rate constant of 4.2 x 103 sec-1 so 37% of the molecules would remain after 240 sec of irradiation.
In a single plane, 16 scans would cause 6-50% bleaching
Material Source:
Pawley: Handbook of Confocal Microscopy

## Slide 27: Förster/Fluorescence Resonance Energy Transfer
Resonance energy transfer can occur when the donor and acceptor molecules are less than 100 Å of one another (preferable 20-50 Å)
Energy transfer is non-radiative which means the donor is not emitting a photon which is absorbed by the acceptor
Förster (Fluorescence) RET (FRET) can be used to spectrally shift the fluorescence emission of a molecular combination.
3rd Ed. Shapiro p 90
4th Ed. Shapiro p 115

## Slide 28: FRET properties
Isolated donor
Donor distance too great
Donor distance correct

## Slide 29: Energy Transfer
Effective between 10-100 Å only
Emission and excitation spectrum must significantly overlap
Donor transfers non-radiatively to the acceptor
PE-Texas Red™
Carboxyfluorescein-Sulforhodamine B
Non radiative energy transfer – a quantum mechanical process of resonance between transition dipoles

## Slide 30: Resonance Energy Transfer
Intensity
Wavelength
Absorbance
DONOR
Absorbance
Fluorescence
Fluorescence
ACCEPTOR
Molecule 1
Molecule 2
Intensity
Molecule 1
Molecule 2
Donor
Acceptor
Fluorescence
Fluorescence

## Slide 31: Bathochromic shift
Bathochromic shift (from Greek βαθύς bathys, "deep"; and χρῶμα chrōma, "color"; hence less common alternate spelling "bathychromic") is a change of spectral band position in the absorption, reflectance, transmittance, or emission spectrum of a molecule to longer wavelength (lower frequency).[1] Because the red color in the visible spectrum has a longer wavelength than most other colors, the effect is also commonly called a red shift.
Hypsochromic shift is a change to shorter wavelength (higher frequency).
https://en.wikipedia.org/wiki/Bathochromic_shift

## Slide 32: Hypsochromic shift
Hypsochromic shift (from ancient Greek ὕψος (upsos) "height"; and χρῶμα chrōma, "color") is a change of spectral band position in the absorption, reflectance, transmittance, or emission spectrum of a molecule to a shorter wavelength (higher frequency). Because the blue color in the visible spectrum has a shorter wavelength than most other colors, this effect is also commonly called a blue shift.
This can occur because of a change in environmental conditions: for example, a change in solvent polarity will result in solvatochromism. A series of structurally related molecules in a substitution series can also show a hypsochromic shift. Hypsochromic shift is a phenomenon seen in molecular spectra, not atomic spectra - it is thus more common to speak of the movement of the peaks in the spectrum rather than lines.
https://en.wikipedia.org/wiki/Hypsochromic_shift

## Slide 33: Hyperchromicity
Hyperchromicity is the increase of absorbance (optical density) of a material. The most famous example is the hyperchromicity of DNA that occurs when the DNA duplex is denatured.[1] The UV absorption is increased when the two single DNA strands are being separated, either by heat or by addition of denaturant or by increasing the pH level. The opposite, a decrease of absorbance is called hypochromicity.
https://en.wikipedia.org/wiki/Hyperchromicity

## Slide 34
In chemistry, solvatochromism is the phenomenon observed when the colour due to a solute is different when that solute dissolved in different solvents.[1][2]
Reichardt's dye dissolved in different solvents
The solvatochromic effect is the way the spectrum of a substance (the solute) varies when the substance is dissolved in a variety of solvents. In this context, the dielectric constant and hydrogen bonding capacity are the most important properties of the solvent. With various solvents there is a different effect on the electronic ground stateand excited state of the solute, so that the size of energy gap between them changes as the solvent changes. This is reflected in the absorption or emission spectrum of the solute as differences in the position, intensity, and shape of the spectroscopic bands. When the spectroscopic band occurs in the visible part of the spectrum solvatochromism is observed as a change of colour. This is illustrated by Reichardt's dye, as shown in the image.
https://en.wikipedia.org/wiki/Solvatochromism

## Slide 35: Measuring Fluorescence Fluorescent Microscope
Dichroic Filter
Objective
Arc Lamp
Emission Filter
Excitation Diaphragm
Ocular
Excitation Filter
EPI-Illumination

## Slide 36: Typical Fluorescence Microscopes
upright
inverted

## Slide 37: Measuring Fluorescence Cameras and emission filters
Color CCD camera does not need optical filters to collect all wavelengths but if you want to collect each emission wavelength optimally, you need a monochrome camera with separate emission filters shown on the right. Alternatives include AOTF or liquid crystal filters.
Camera goes here

## Slide 38
Stained frozen C57BL/6 mouse spleen tissue using antibodies against CD4 and CD8a to detect T cells, B220 to stain B cells, and CD169 and F4/80 to detect tissue-resident macrophages. For this staining, we took advantage of antibodies directly-conjugated to bright photostable fluorophores including the Brilliant Violet™ and Alexa Fluor® dyes.
https://www.biolegend.com/microscopy
Human paraffin-embedded colon tissue slices were Alexa Fluor® 594 anti-human Galectin-9 (clone 9M1-3, red) and Alexa Fluor® 647 anti-Vimentin (clone O91D3, green) and DAPI (blue).
https://www.biolegend.com/microscopy

## Slide 39: Probes for Proteins
FITC 488 525
PE 488 575
APC 630 650
PerCP™ 488 680
Cascade Blue 360 450
Coumerin-phalloidin 350 450
Texas Red™ 610 630
Tetramethylrhodamine-amines 550 575
CY3 (indotrimethinecyanines) 540 575
CY5 (indopentamethinecyanines) 640 670
Probe Excitation Emission

## Slide 40: Additional dyes
https://www.biolegend.com/media_assets/brilliantviolet/BVSpectraChart062512.jpg

## Slide 41: Panel Selector tools
e.g. https://www.biolegend.com/panelselector

## Slide 42: Probes for Nucleic Acids
Hoechst 33342 (AT rich) (uv) 346 460
DAPI (uv) 359 461
POPO-1 434 456
YOYO-1 491 509
Acridine Orange (RNA) 460 650
Acridine Orange (DNA) 502 536
Thiazole Orange (vis) 509 525
TOTO-1 514 533
Ethidium Bromide 526 604
PI (uv/vis) 536 620
7-Aminoactinomycin D (7AAD) 555 655

## Slide 43: DNA Probes
AO
  Metachromatic dye
    concentration dependent emission
    double stranded NA - Green
    single stranded NA - Red
AT/GC binding dyes
  AT rich: DAPI, Hoechst, quinacrine
  GC rich: antibiotics bleomycin, chromamycin A3, mithramycin, olivomycin, rhodamine 800

## Slide 44: Probes for Ions
INDO-1 Ex350 Em405/480
QUIN-2 Ex350 Em490
Fluo-3 Ex488 Em525
Fura -2 Ex330/360 Em510
INDO-1: 1H-Indole-6-carboxylic acid, 2-[4-[bis[2-[(acetyloxy)methoxy]-2- oxoethyl]amino]-3-[2-[2-[bis[2- [(acetyloxy)methoxy]-2-oxoetyl]amino]-5- methylphenoxy]ethoxy]phenyl]-, (acetyloxy)methyl ester [C47H51N3O22 ] (just in case you want to know….!!)
Indo-1
FLUO-3: Glycine, N-[4-[6-[(acetyloxy)methoxy]-2,7- dichloro-3-oxo-3H-xanthen-9-yl]-2-[2-[2- [bis[2-[(acetyloxy)methoxy]-2- oxyethyl]amino]-5- methylphenoxy]ethoxy]phenyl]-N-[2- [(acetyloxy)methoxy]-2-oxyethyl]-, (acetyloxy)methyl ester

## Slide 45: pH Sensitive Indicators
SNARF-1 488 575
BCECF 488 525/620
440/488 525
Probe Excitation Emission
SNARF-1: Benzenedicarboxylic acid, 2(or 4)-[10-(dimethylamino)-3-oxo-3H- benzo[c]xanthene-7-yl]-
BCECF: Spiro(isobenzofuran-1(3H),9'-(9H) xanthene)-2',7'-dipropanoic acid, ar-carboxy-3',6'-dihydroxy-3-oxo-
C27H20O11
C27H19NO6

## Slide 46: Probes for Oxidation States
DCFH-DA (H2O2) 488 525
HE (O2-) 488 590
DHR 123 (H2O2) 488 525
Probe Oxidant Excitation Emission
DCFH-DA - dichlorofluorescin diacetate
HE - hydroethidine 3,8-Phenanthridinediamine, 5-ethyl-5,6-dihydro-6-phenyl-
DHR-123 - dihydrorhodamine 123 Benzoic acid, 2-(3,6-diamino-9H-xanthene-9-yl)-, methyl ester
DCFH-DA: 2',7'-dichlorodihydrofluorescein diacetate (2',7'-dichlorofluorescin diacetate; H2DCFDA)
C24H16Cl2O7
C21H21N3
C21H18N2O3

## Slide 47: Specific Organelle Probes
BODIPY Golgi 505 511
NBD Golgi 488 525
DPH Lipid 350 420
TMA-DPH Lipid 350 420
Rhodamine 123 Mitochondria 488 525
DiO Lipid 488 500
diI-Cn-(5) Lipid 550 565
diO-Cn-(3) Lipid 488 500
Probe Site Excitation Emission
BODIPY - borate-dipyrromethene complexes NBD - nitrobenzoxadiazole
DPH – diphenylhexatriene TMA - trimethylammonium

## Slide 48: Other Probes of Interest
GFP - Green Fluorescent Protein
  GFP is from the chemiluminescent jellyfish Aequorea victoria
  excitation maxima at 395 and 470 nm (quantum efficiency is 0.8) Peak emission at 509 nm
  contains a p-hydroxybenzylidene-imidazolone chromophore generated by oxidation of the Ser-Tyr-Gly at positions 65-67 of the primary sequence
  Major application is as a reporter gene for assay of promoter activity
  requires no added substrates
Note: 2008 Nobel prize for Chemistry was for GFP (Roger Tsien)
Roger Tsien and J.Paul Robinson in Rome in March 2009 when Roger was awarded the Romulus and Remus award by the Mayor of Rome

## Slide 49: Multiple Emissions
Many possibilities for using multiple probes with a single excitation
Multiple excitation lines are possible
Combination of multiple excitation lines or probes that have same excitation and quite different emissions
  e.g. Calcein AM and Ethidium (ex 488 nm)
  emissions 530 nm and 617 nm

## Slide 50: Filter combinations
The band width of the filter will change the intensity of the measurement

## Slide 51: Fluorescence Overlap

## Slide 52: Fluorescence Overlap

## Slide 53: Fluorescence Overlap
a
b
Overlap of FITC fluorescence in PE PMT
Overlap of PE fluorescence in FITC PMT
This is your bandpass filter

## Slide 54: Fluorescence
The longer the wavelength the lower the energy
The shorter the wavelength the higher the energy
  eg. UV light from sun - this causes the sunburn, not the red visible light
The spectrum is independent of precise excitation line but the intensity of emission is not

## Slide 55: Mixing fluorochromes
When there are two molecules with different absorption spectra, it is important to consider where a fixed wavelength excitation should be placed. It is possible to increase or decrease the sensitivity of one molecule or another.

## Slide 56: Mixing fluorochromes
When there are two molecules with different absorption spectra, it is important to consider where a fixed wavelength excitation should be placed. It is possible to increase or decrease the sensitivity of one molecule or another.

## Slide 57
J. Paul Robinson, Class lecture notes, BMS 631
Excitation of 3 Dyes with emission spectra

## Slide 58: Change of Excitation
J. Paul Robinson, Class lecture notes, BMS 631

## Slide 59: LEDs for microscopy
https://www.photonics.com/Articles/LEDs_in_Microscopy_An_Emerging_Research_Tool/a62272
Figure 3. Fluorescent bovine pulmonary artery endothelial cells. The cell nucleus is shown blue, the green is tubulin and the red shows actin. Courtesy of Jordi Recasens/Izasa Scientific.

## Slide 60: Commercial Microscopes use LEDs
https://www.thermofisher.com/us/en/home/life-science/cell-analysis/cellular-imaging/cell-imaging-systems/floid-cell-imaging-station.html

## Slide 61: LEDs used in Microscopy
https://www.photonics.com/Articles/LEDs_in_Microscopy_An_Emerging_Research_Tool/a62272
Figure 4. This graph shows a 385-nm LED and Hg spectrum overlaid on DAPI absorption and emission, with an example excitation and emission filter. Courtesy of Gerard Whoriskey/CoolLED Ltd.
340-nm LEDs are also available for deep UYV probes like Fura 2

## Slide 62: Conclusions
Fluorescence is the primary energy source for confocal microscopes
Dye molecules must be close to, but below saturation levels for optimum emission
Fluorescence emission is longer than the exciting wavelength
The energy of the light increases with reduction of wavelength
Fluorescence probes must be appropriate for the excitation source and the sample of interest
Correct optical filters must be used for multiple color fluorescence emission