## Slide 1: BMS 524 - “Introduction to Confocal Microscopy and Image Analysis”
Lecture 9: Components of the microscope
Department of Basic Medical Sciences,
School of Veterinary Medicine
Weldon School of Biomedical Engineering
Purdue University
J. Paul Robinson, Ph.D.
Distinguished Professor of Cytometry
& Professor of Biomedical Engineering
Director, Purdue University Cytometry Laboratories, Purdue University
This lecture was last updated in March, 2017
Find other PUCL Educational Materials at http://www.cyto.purdue.edu/class
These slides are intended for use in a lecture series. Copies of the slides are distributed and students encouraged to take their notes on these graphics. All material copyright J.Paul Robinson unless otherwise stated. No reproduction of this material is permitted without the written permission of J. Paul Robinson. Except that our materials may be used in not-for-profit educational institutions ith appropriate acknowledgement. It is illegal to upload this lecture to CourseHero or any other site.
See also Microscopy Encyclopedia chapter
http://www.cyto.purdue.edu/flowcyt/research/pdfs/micro.pdf
See more lectures and slides at:
http://www.cyto.purdue.edu/flowcyt/educate/pptslide.htm

## Slide 2: Learning Objectives
Understand:
Components of a confocal microscope system
Optical pathways used in systems
Optical resolution - Airy disks
Reflected light/backscatter imaging

## Slide 3
http://zeiss-campus.magnet.fsu.edu/articles/superresolution/introduction.html

## Slide 4: Benefits of Confocal Microscopy
Reduced blurring of the image from light scattering
Increased effective resolution
Improved signal to noise ratio
Clear examination of thick specimens
Z-axis scanning
Depth perception in Z-sectioned images
Magnification can be adjusted electronically

## Slide 5: Fluorescent Microscope
Objective
Arc Lamp
Emission Filter
Excitation Diaphragm
Ocular
Excitation Filter

## Slide 6: Confocal Principle
Objective
Laser
Emission Pinhole
Excitation Pinhole
PMT
Emission
Filter
Excitation Filter

## Slide 7: Confocal microscopy
Figure from D. Andrews, I. S. Harper and J. R. Swedlow

## Slide 8
Fluorescent Microscope
Objective
Arc Lamp
Emission Filter
Excitation Diaphragm
Ocular
Excitation Filter
Objective
Laser
Emission Pinhole
Excitation Pinhole
PMT
Emission
Filter
Excitation Filter
Confocal Microscope

## Slide 9: Wide-field vs confocal vs 2-p
Drawing by P. D. Andrews, I. S. Harper and J. R. Swedlow

## Slide 10: Bio-Rad MRC 1024 (Circa 1993-2000) (UV 350 nm microscopy) The first fully UV (353nm) confocal microscope – thus the protection canopy around the microscope
Historical Slide

## Slide 11: Radiance 2100MP in Purdue University Cytometry Laboratories
Zeiss Confocal Scopes
Nikon
Leica

## Slide 12: Original Confocal layout using large lasers
UV Laser
Kr-Ar Laser
Optical Mixer
Scanhead
Microscope
Computer
analysis
Lasers
Optical
Components
Microscope
Computer

## Slide 13
e.g. MRC 1024 System
Light Path Concept for Confocal
PMT

## Slide 14: Optical Mixer - MRC 1024 UV
Argon Laser
Argon-
Krypton
Laser
Fast Shutter
UV Correction
Optics
Filter
Wheels
To Scanhead
UV
Visible
353,361 nm
488, 514 nm
488,568,647 nm
Beam Expander
Excitation optics

## Slide 15: MRC 1024 Scanhead
From Laser
To and from Scope
3
2
1
PMT
Galvanometers
Emission
Filter
Wheel

## Slide 16: Scanning Galvanometers
x
y
Laser in
Laser out
Point Scanning
To
Microscope

## Slide 17: MRC 1024 Scanhead
From Laser
To and from Scope
3
2
1
PMT
Galvanometers
Emission
Filter
Wheels
This dichroic separates the excitation from the emission signals
PMT
PMT

## Slide 18: Confocal problems
Destroyed optical excitation filters using high power lasers
Hole burned in filter

## Slide 19: Zeiss 710 system
Materials from Zeiss publication

## Slide 20: Zeiss 710
Materials from Zeiss publication

## Slide 21: Complex optical pathways
BioRads “Twingate” beam splitters from early 1990s

## Slide 22: Zeiss 710
Materials from Zeiss publication

## Slide 23: Zeiss 710 spectral system
http://zeiss-campus.magnet.fsu.edu/tutorials/spectralimaging/quasar34ch/index.html

## Slide 24: Zeiss 710 spectral system
Images from: http://zeiss-campus.magnet.fsu.edu/tutorials/spectralimaging/quasar34ch/index.html

## Slide 25: Spectral CLSM (Zeiss)
Zeiss LSM 5 LIVE
A new system concept of optics and electronics
Dispersion grating
Multianode PMT
Materials from leica publication

## Slide 26
http://www.zeiss.com/C12567BE00472A5C/EmbedTitelIntern/Article_Nature_Methods/$File/Faster_than_real-time.pdf
512 x 50 pixels – 1010 frames/sec
512x512 pixels – 100 FPS

## Slide 27: Spectral CLSM (Leica)
Original Leica System
Fluorescence signal was dispersed on a prism, to provide spectral analysis capability
Excitation via acousto-optical
tunable filter:
Free selection of excitation lines
Individual line attenuation
Less crosstalk
Less fading
Materials from leica publication

## Slide 28
To Scanhead
From
Scanhead
Review of the fundamental principles

## Slide 29: The Scan Path of the Laser Beam
767, 1023, 1279, etc
511, 1023
0
0
Start
Specimen
Frames/Sec # Lines
1 512
2 256
4 128
8 64
16 32
Early confocal systems
Leica SP2
400 Hz
512 line image takes 1.28 seconds
4096 line image takes 10.24 seconds
Nikon A1 has up to 4096 lines
Zeiss 710 has up to 6000 lines

## Slide 30: How a Confocal Image is Formed
Condenser
Lens
Pinhole 1
Pinhole 2
Objective
Lens
Specimen
Detector
Modified from: Handbook of Biological Confocal Microscopy. J.B.Pawley, Plennum Press, 1989

## Slide 31: Fundamental Limitations of Confocal Microscopy
From
Source
To Detector
.
x,y,z
2
n2 photons
2
1
n1 photons
1
z
y
x
VOXEL
3D space
PIXEL
2D space
From: Handbook of Biological Confocal Microscopy. J.B.Pawley, Plennum Press, 1989
Resolutionx,y = λ / 2[η • sin(α)]
Resolutionz = 2λ / [η • sin(α)]2
(NA)
Is around 150 nanometers in the lateral dimension and approaching 400 nanometers in the axial dimension

## Slide 32: Super Resolution Microscopy
In super-resolution microscopy both lateral and axial resolution can be measured in the tens of nanometers. Thus they are able to resolve features beneath the diffraction limit by switching fluorophores on and off sequentially in time so that the signals can be recorded consecutively.
STED (Stimulated Emission Depletion; from the Stefan Hell laboratory)
SSIM (Saturated Structured Illumination Microscopy; pioneered by Mats Gustafsson).
PALM (PhotoActivated localization Microscopy; introduced by Eric Betzig and Harald Hess)
STORM (Stochastic Optical Reconstruction Microscopy; first reported by Xiaowei Zhang).
(to be discussed in a later lecture in detail)

## Slide 33: Microscopy Definitions
Near-field, (NSOM) the specimen is imaged within a region having a radius much shorter than the illumination wavelength.
Far-field, the specimen many thousands of wavelengths away from the objective (often a millimeter or more) and is limited in resolution by diffraction of the optical wave-fronts as they pass through the objective rear aperture (this would include brightfield, DIC, and phase contrast, widefield fluorescence, confocal, and multiphoton are considered far-field, diffraction-limited instruments.

## Slide 34: I5M and 4Pi Microscopy
http://zeiss-campus.magnet.fsu.edu/articles/superresolution/introduction.html

## Slide 35: Light Sheet Microscopy
https://www.researchgate.net/publication/306283216_New_Methods_to_Improve_Large-Scale_Microscopy_Image_Analysis_with_Prior_Knowledge_and_Uncertainty/figures?lo=1
Gualda et al, Cytometry Part A 91A: 144151, 2017
Light sheet flow cytometry

## Slide 36: Gray Level & Pixilation (reminder)
Analogous to intensity range
For computer images each pixel is assigned a value. If the image is 8 bit, there are 28 or 256 levels of intensity If the image is 10 bit there are 1024 levels, 12 bit 4096 levels etc.
The intensity analogue of a pixel is its grey level which shows up as brightness.
The display will determine the possible resolution since on a TV screen, the image can only be displayed based upon the number of elements in the display. Of course, it is not possible to increase the resolution of an image by attributing more “pixels” to it than were collected in the original collection!

## Slide 37: Sampling Theory (Reminder)
The Nyquist Theorem
  Nyquest theory describes the sampling frequency (f) required to represent the true identity of the sample.
  i.e., how many times must you sample an image to know that your sample truly represents the image?
  In other words to capture the periodic components of frequency f in a signal we need to sample at least 2f times
Nyquist claimed that the rate was 2f. It has been determined that in reality the rate is 2.3f - in essence you must sample at least 2 times the highest frequency.
For example in audio, to capture the 22 kHz in the digitized signal, we need to sample at least 44.1 kHz (unless of course you can’t hear 22k Hz and then you don’t need 44.1 kHz!!!!)
(See previous lecture on Nyquist Theorem)

## Slide 38: Digital Zoom
1 x
1024 points
2 x
1024 points
4 x
1024 points
Note that we have reduced the field of view of the sample
Note #2: There will only be a single zoom value where
optimal resolution can be collected.

## Slide 39: 3D imaging using CLSM
Videos made on Biorad 1024 Confocal in Robinson Lab, Purdue ~1996-1999

## Slide 40: Live Human Neutrophils
Phagocytosing yeast
Reference: Bassoe-C-F, Nianyu Li, Kathy Ragheb, Gretchen Lawler, Jennie Sturgis, J. Paul Robinson, Cytometric Investigations of Phagosomes, Mitochondria and Acidic Granules in Human neutrophils. Cytometry 2003:51B:21-29.

## Slide 41: Cells in ECM
Videos made on Biorad 1024 Confocal in Robinson Lab, Purdue ~1996-1999
Studies with Dr. Sherry Harbin, Purdue BME

## Slide 42: Reflection Imaging
Backscattered light imaging
Same wavelength as excitation
Advantages: no photobleaching since not using a photo-probe (note: does not mean no possible damage to specimen)
Problems: optical reflections from components of microscope

## Slide 43: CLSM can utilize autofluorescence: Example - topography of SIS
Raw image
Deconvolved image
/Users/wombat/Syncplicity/PI4D image symposium materials/sis_af_02.mp4
Video made on Biorad 1024 Confocal in Robinson Lab, Purdue

## Slide 44: Backscattered light imaging – another CLSM modality
Collagen gel (ECM model) visualized by BSL-confocal microscopy. Volume- (left) and surface-rendering (right) of a confocal dataset.
Collagen gel – an optical section
Movies/gel_bsl_02.avi
Movies/Gel_bs2.mpg
Videos made on Biorad 1024 Confocal in Robinson Lab, Purdue

## Slide 45: Backscattered light: tissue imaging
Left: Optical sections of SIS visualized with BSL-confocal technique. Bottom: color-coded height map revealing the topography of height map of SIS

## Slide 46: Example: BSL and AF signals combined
HepG2 cells grow embedded within a collagen matrix (animation)
Videos made on Biorad 1024 Confocal in Robinson Lab, Purdue ~1996-1999

## Slide 47: Reflected light images
CD-ROM pits imaged on a 1024 Bio-Rad confocal
Collagen imaged by
a 1024 Confocal
Videos made on Biorad 1024 Confocal in Robinson Lab, Purdue ~1996-1999

## Slide 48: Backscattered light and autofluorescence signals combined collagen gel & HepG2 cells
Videos made on Biorad 1024 Confocal in Robinson Lab, Purdue ~1996-1999

## Slide 49: Imaging spectroscopy

## Slide 50: Issues for good confocal imaging
Axial Resolution
  Must determine the FWHM (full width half maximum) intensity values of a vertical section of beads
Field Flatness
  Must be able to collect a flat field image over a specimen - or z-axis information will be inaccurate
Chromatic Aberration
  must test across an entire field that emission is constant and not collecting radial or tangential artifacts due to chromatic aberration in objectives
Z-drive precision and accuracy
  must be able to reproducibility measure distance through a specimen - tenths of microns will make a big difference over 50 microns

## Slide 51: SUMMARY SLIDE
Components of a confocal system
Optical pathways
Optical resolution
Sampling rate (Nyquist Theorem)
Reflection imaging
Current imaging systems