## Slide 1: Live Cell Imaging Applications in Confocal Microscopy BMS 524 - “Introduction to Confocal Microscopy and Image Analysis” Purdue University Department of Basic Medical Sciences, School of Veterinary Medicine
UPDATED January 2023
J. Paul Robinson, Ph.D.
SVM Professor of Cytomics
Professor of Biomedical Engineering
Director, Purdue University Cytometry Laboratories
These slides are intended for use in a lecture series. Copies of the graphics are distributed and students encouraged to take their notes on these graphics. All material copyright J. Paul Robinson unless otherwise stated, however, the material may be freely used for lectures, tutorials and workshops. It may not be used for any commercial purpose including uploading to CourseHero.
The text for this course is Pawley “Introduction to Confocal Microscopy”, Plenum Press, 2nd Ed. A number of the ideas and figures in these lecture notes are taken from this text.

## Slide 2: Learning Objectives
Identify applications of imaging related to live cell work
Identify tools useful for doing live cell work
Identify reagents compatible with live cell imaging
Identify problems associated with live cell imaging
Show some examples of live cell imaging

## Slide 3: Applications
Organelle Structure
Probe ratioing
Conjugated antibodies
DNA/RNA
Cytochemical Identification
Oxidative Metabolism
Exotic Applications

## Slide 4: Applications
Organelle Structure & Function
  Mitochondria (Rhodamine 123)
  Golgi (C6-NBD-Ceramide)
  Actin (NBD-Phaloidin)
  Lipid (DPH)

## Slide 5
Step 1: Cell Culture
Step 2: Cell Wash
Lab-Tek
1
2
3
4
5
6
7
8
top view
side view
170 M coverslip
Step 3: Transfer to Lab-Tek plates
confocal microscope
oil immersion
objective
37o heated stage
stimulant/inhibitor added
Step 4: Addition of DCFH-DA, Indo-1, or HE
Below: the culture dishes for live cell imaging using a confocal microscope and high NA objectives.

## Slide 6: Lumascope
The LumaScope is a low price, compact, USB-based, inverted fluorescence microscope. The system fits on a shelf inside an incubator or inside a hood. It has a built-in camera and light source. This runs on a laptop/desktop using just one USB cable, captures fluorescence or brightfield images and has options of 2.5x, 4x, 10x, 20x, 40x and 100x (oil) lenses. The software allows to do time-laps study. The machine does not have a mechanical shutter to block the light (to avoid photo-toxicity or photo-bleaching), but the software can control the light to turn on/off to less expose the cells to the light.

## Slide 7: Nikon live cell imaging
Image source: http://www.microscopyu.com/articles/livecellimaging/livecellmaintenance.html

## Slide 8: Live Cell Chambers
http://www.microscopyu.com/articles/livecellimaging/culturechambers.html

## Slide 9: Examples of micro-incubation systems
Images from: http://www.harvardapparatus.com/webapp/wcs/stores/servlet/haisku3_10001_11051_68114_-1_HAI_ProductDetail_N_37474_37500_44260

## Slide 10
Confocal System
Culture System
Photos taken in Purdue University Cytometry Labs
Photo taken from Nikon promotion material

## Slide 11: Live cell manipulation devices
Image source: http://www.warneronline.com/products.cfm?CFID=6704527&CFTOKEN=47754450

## Slide 12: LabtekTM culture chambers
http://www.microscopyu.com/articles/livecellimaging/culturechambers.html

## Slide 13: Useful cell lines
Table from: http://www.microscopyu.com/articles/livecellimaging/livecellmaintenance.html

## Slide 14: Growth conditions
Table from: http://www.microscopyu.com/articles/livecellimaging/livecellmaintenance.html

## Slide 15: Example of DIC and Fluorescence
Human cheek epithelial cells (from JPR!) stained with Hoechst 33342 - wet prep, 20 x objective, 3 x zoom (Bio-Rad 1024 MRC) (Image from JPR lab)
Giardia (DIC image)
(no fluorescence)
(photo taken from a 35 mm slide and scanned - cells were live when photographed)
(JPR lab)

## Slide 16
Fluorescence Microscope image of Hoechst stained cells (plus DIC)
Image collected with a 470T Optronics cooled camera (Image from JPR lab)

## Slide 17: Measurement of DNA
Use for DNA content and cell viability
  33342 for viability
Less needed to stain for DNA content than for viability
  decrease nonspecific fluorescence
Low laser power decreases CVs
G0-G1
S
G2-M
Fluorescence Intensity
# of Events
Notes: Ratio of 33342 fluorescence
510-560 (green) vs. 390-440 (violet)
different for different subclasses of chicken thymocytes
shift of 33258 emission toward green as [dye] increased
fixed rat thymocytes
low dye/Protein ratio, Hi affinity binding with hi QE yield blue fluorescence
as increased [dye], lower affinity binding to secondary sites, emission shifts toward green, and DNA precipitates

## Slide 18: PI - Cell Viability
How the assay works:
PI cannot normally cross the cell membrane
If the PI penetrates the cell membrane, it is assumed to be damaged
Cells that are brightly fluorescent with the PI are damaged or dead
PI
PI
PI
PI
PI
PI
PI
PI
PI
PI
PI
PI
PI
PI
Viable Cell Damaged Cell

## Slide 19
Flow cytometric scatter plot of gamma irradiated C. parvum oocysts. The oocysts region is clearly distinguished from ghosts and debris. Images on the right show Sytox green fluorescence and transmission images of these regions. Note ghosts do not take up Sytox green dye.
Fluorescence Transmission
Crypto
10
0
10
1
10
2
10
3
10
4
10
0
10
1
10
2
10
3
10
4
Green Fluorescence
Forward Scatter
Side Scatter
Flow Cytometry Dot Plot
oocysts
debris
ghosts
Notes: Crypto isolated from cow (purchased)
gamma irradiated 900 Krads (kills alm ost all)
Stained with Sytox Green (membrane impermeant) i.e. gamma irradiation damages oocyst wall to all dye to penetrate and stain DNA
Ghosts - don’t stain because trophozoites have left and there is no DNA to stain
Dead cells stain because of the damaged membrane from the gamma irradiation
Live cells will not stain either

## Slide 20: Specific Organelle Probes
BODIPY Golgi 505 511
NBD Golgi 488 525
DPH Lipid 350 420
TMA-DPH Lipid 350 420
Rhodamine 123 Mitochondria 488 525
DiO Lipid 488 500
diI-Cn-(5) Lipid 550 565
diO-Cn-(3) Lipid 488 500
Probe Site Excitation Emission
BODIPY - borate-dipyrromethene complexes
NBD - nitrobenzoxadiazole
DPH - diphenylhexatriene
TMA - trimethylammonium

## Slide 21: Organelle Function
Mitochondria Rhodamine 123
Endosomes Ceramides
Golgi BODIPY-Ceramide
Endoplasmic Reticulum DiOC6(3) Carbocyanine

## Slide 22: Calcium Related Applications
Probe Ratioing
  Calcium Flux (Indo-1)
  pH indicators (BCECF, SNARF)
Molecule-probe Excitation Emission
Calcium - Indo-1 351 nm 405, >460 nm
Calcium- Fluo-3 488 nm 525 nm
Calcium - Fura-2 363 nm >500 nm
Calcium - Calcium Green 488 nm 515 nm
Magnesium - Mag-Indo-1 351 nm 405, >460 nm
Phospholipase A- Acyl Pyrene 351 nm 405, >460 nm

## Slide 23: Probes for Ions
INDO-1 Ex350 Em405/480
QUIN-2 Ex350 Em490
Fluo-3 Ex488 Em525
Fura -2 Ex330/360 Em510

## Slide 24: Ionic Flux Determinations
Calcium Indo-1
Intracellular pH BCECF
How the assay works:
Fluorescent probes such as Indo-1 are able to bind to calcium in a ratiometric manner
The emission wavelength decreases as the probe binds available calcium
Time (Seconds)
0
36
72
108
144
180
Stimulation
0
0.1
0.2
0.3
0.4
0.5
0.6
0.7
0.8
0
50
100
150
200
Ratio: intensity of 460nm / 405nm signals
Time (seconds)
Flow Cytometry Image Analysis

## Slide 25: Calcium Flux
0
0.1
0.2
0.3
0.4
0.5
0.6
0.7
0.8
0
50
100
150
200
Ratio: intensity of 460nm / 405nm signals
Time (seconds)
Time (Seconds)
0
36
72
108
144
180
Stimulation
Flow Cytometry
Image Cytometry

## Slide 26: Oxidative-Related Reactions
Superoxide Hydroethidine
Hydrogen Peroxide Dichlorofluorescein
Glutathione levels Monobromobimane
Nitric Oxide Dichlorofluorescein
Nitric Oxide DAF-FM
Nitric Oxide Indicators: DAF-FM and DAF-FM Diacetate
D-23841 DAF-FM (4-amino-5-methylamino- 2',7'-difluorescein)
D-23842 DAF-FM diacetate (4-amino-5-methylamino- 2',7'-difluorofluorescein diacetate)
D-23844 DAF-FM diacetate (4-amino-5-methylamino- 2',7'-difluorofluorescein diacetate) *special packaging*
(Data from Invitrogen website)

## Slide 27: DAF-FM DAF-FM Diacetate (4-Amino-5-Methylamino-2',7'-Difluorofluorescin Diacetate)
DAF-FM is not a reversible equilibrium sensor, limiting its ability to track rapid fluctuations of the target analyte (NO) in real time.
DAF-FM is a reagent that is used to detect and quantify low concentrations of nitric oxide (NO). It is essentially nonfluorescent until it reacts with NO to form a fluorescent benzotriazole. DAF-FM fluorescence can be detected by any instrument that can detect fluorescein, including flow cytometers, microscopes, fluorescent microplate readers and fluorometers.
Ex/Em of DAF-FM: ~495/515 nm
The fluorescence quantum yield of DAF-FM is ~0.005, but increases about 160-fold, to ~0.81, after reacting with NO
Data and Image from http://products.invitrogen.com/ivgn/product/D23842

## Slide 28
DCFH-DA DCFH DCF
COOH
H
Cl
O
O-C-CH3
O
CH3-C-O
Cl
O
COOH
H
Cl
OH
HO
Cl
O
COOH
H
Cl
O
HO
Cl
O
Fluorescent
Hydrolysis
Oxidation
2’,7’-dichlorofluorescin
2’,7’-dichlorofluorescin diacetate
2’,7’-dichlorofluorescein
Cellular Esterases
H2O2
DCFH-DA
DCFH-DA
DCFH
DCF
H O
2 2
Lymphocytes
Monocytes
Neutrophils
log FITC Fluorescence
.1
1000
100
10
1
0
20
40
60
counts
PMA-stimulated PMN
Control
80

## Slide 29: Hydroethidine
HE Ethidium
N
CH2CH3
NH2
H2N
H
Br-
N
CH2CH3
NH2
H2N
+
O2-
Phagocytic Vacuole
SOD
H2O2
NADPH
NADP
O2
NADPH Oxidase
OH-
O2-
DCF
HE
O2-
H2O2
DCF
Example: Neutrophil Oxidative Burst

## Slide 30: Macrovascular Endothelial Cells in Culture
Time (minutes)
0
60
hydroethidine

## Slide 31
Hydrogen peroxide measurements with DCFH-DA
0
200
400
600
800
1000
1200
1400
1600
1800
2000
0
500
1000
1500
2000
2500
3000
Time in seconds
cell 1
cell 2
cell 3
cell 4
cell 5
% change (DCF fluorescence)
525 nm
1
2
3
4
5
Step 6B: Export data from measured
regions to Microsoft Excel
Step 7B: Export data from Excel data
base to Delta Graph
Change in fluorescence was measured using Bio-Rad software and the data exported to a spread sheet for analysis.

## Slide 32
Superoxide measured with hydroethidine
Export data from Excel data
base to Delta Graph
Export data from measured
regions to Microsoft Excel
cell 1
cell 2
cell 3
cell 4
cell 5
Change in fluorescence was measured using Bio-Rad software and the data exported to a spread sheet for analysis.
%change (DCF fluorescence)
-200
0
200
400
600
800
1000
1200
1400
1600
1800
cell 1
cell 2
cell 3
cell 4
cell 5
Time in seconds
1000 1200 1400 1600 1800
600 800
200 400

## Slide 33: H2O2 stimulation and DCF & Ethidium loading in Rat Pulmonary Artery Endothelial Cells
ENDO HBSS
ENDO HBSS TNFa
ENDO L-arg
ENDO/ L-arg TNFa
ENDO/ D-arg
ENDO/ D-arg TNFa
Endo + 200uM H2O2
Endo + 200uM H2O2
Endo + 200uM H2O2
Endo / TNFa + 200uM H2O2
Endo / TNFa + 200uM H2O2
Endo / TNFa + 200uM H2O2
Endo / L-arg + 200uM H2O2
Endo / L-arg + 200uM H2O2
Endo / L-arg + 200uM H2O2
Endo / L-arg TNFa + 200uM H2O2
Endo / L-arg TNFa + 200uM H2O2
Endo / L-arg TNFa + 200uM H2O2
Endo / D-arg + 200uM H2O2
Endo / D-arg + 200uM H2O2
Endo / D-arg + 200uM H2O2
Endo / D-arg TNFa + 200uM H2O2
Endo / D-arg TNFa + 200uM H2O2
Endo / D-arg TNFa + 200uM H2O2
0
20
40
60
80
100
120
140
160
180
200
0
20
40
60
80
100
120
140
Time (minutes)
Mean EB Fluorescence
.
200uM
H2O2
added
Time (seconds)
DCF Fluorescence
Confocal System - Fluorescence Measurements
200uM
H2O2
added
24 treatments - 5000 cells each

## Slide 34: pH Sensitive Indicators
SNARF-1 488 575
BCECF 488 525/620
440/488 525
[2’,7’-bis-(carboxyethyl)-5,6-carboxyfluorescein]
Probe Excitation Emission

## Slide 35: Exotic Applications of Confocal Microscopy
FRAP (Fluorescence Recovery After Photobleaching)
Release of “Caged” compounds
Lipid Peroxidation (Parinaric Acid) Difficult to do with confocal, but possible with 2P (excitation is 325 nm)
Membrane Fluidity (DPH)

## Slide 36: “Caged” Photoactivatable Probes
Ca++: Nitr-5
Ca++ - buffering: Diazo-2
IP3
cAMP
cGMP
ATP
ATP--S
Available Probes
Principle: Nitrophenyl blocking groups e.g. nitrophenyl ethyl ester undergoes photolysis upon exposure to UV light at 340-350 nm

## Slide 37: Release of “Caged” Compounds
UV Beam
Release of “Cage”
Culture dish

## Slide 38: Release of Caged Compounds
Time (seconds) after UV FLASH
Release of Caged Nitric Oxide in
Attached PMN
0
50
100
150
200
250
0
20
40
60
80
100
120
140
160
Fluorescence Emission at 515 nm
C
D
UV excited
Control Region
Time (seconds) CONTROL
0
50
100
150
200
250
CONTROL STUDY
Fluorescence Emission at 515 nm
0
100
200
300
400

## Slide 39: Membrane Polarization
Polarization/fluidity Diphenylhexatriene
How the assay works: The DPH partitions into liphophilic portions of the cell and is excited by a polarized UV light source. Polarized emissions are collected and changes can be observed kinetically as cells are activated.
An image showing DPH fluorescence in cultured endothelial cells.

## Slide 40
1
2
3
3
2
1
405/35 nm
460 nm
Calcium ratios with Indo-1
Changes in the fluorescence were measured using the Bio-Rad calcium ratioing software. The same region in each wave length was measured and the relative change in each region was recorded and exported to a spread sheet for analysis..
Export data from measured regions to Microsoft Excel
Export data from Excel data
base to Delta Graph
50
100
150
200
0
0.1
0.2
0.3
0.4
0.5
0.6
0.7
0.8
0
cell 1
cell 2
cell 3
Ratio: intensity1 (460nm) / intensity2 (405/35nm)

## Slide 41: FRAP
Intense laser Beam
Bleaches Fluorescence
Recovery of fluorescence
10 seconds
30 seconds
Zero time
Time
%F

## Slide 42: 4D confocal microscopy
Time vs 3D sections
Used when evaluating kinetic changes in tissue or cells
Requires fast 3D sectioning
Difficult to evaluate

## Slide 43: 4D Imaging
Time
1 2 3 4 5
This could also be achieved using an X-Z scan on a point scanner or
Something like the Zeiss Live instrument which is very fast..

## Slide 44: Imaging 3D ECM structures
Mainly collagen-based materials
Usually 40-120 microns thick
Require both transmitted and fluorescent signals
Often require significant image processing to extract information

## Slide 45
Work done with Prof Sherry Harbin

## Slide 46
Work done with
Prof Sherry Harbin

## Slide 47: Other functional confocal studies
2P kidney imaging (Bio-rad 2100-PUCL)
Confocal – epithelial cells (Biorad 1024)
Cheek epithelial cell (PUCL)

## Slide 48
Bassoe-C-F, Nianyu Li, Kathy Ragheb, Gretchen Lawler, Jennie Sturgis, J. Paul Robinson, Cytometric Investigations of Phagosomes, Mitochondria and Acidic Granules in Human neutrophils. Cytometry 2003 Jan;51B(1):21-9.
Cytometry 2003 Jan;51B(1):21-9.
(PUCL)

## Slide 49: Live bacterial analysis (Robinson lab) (PUCL)

## Slide 50: Pinecone spores (PUCL)
Cork
(PUCL)

## Slide 51: Thick Tissue - Bone and Cartilage
Very difficult to image thick specimens
Can use live specimens if appropriately stained
Special preparation techniques

## Slide 52: Commercial Systems
https://www.keyence.com/landing/lpc/live-cell-imaging.jsp?aw=google-kaenBZ214101bb&k_clickid=de902564-219e-44df-84bb-34e61c4c3f2f&gclid=CjwKCAiA_P3jBRAqEiwAZyWWaHDcrVmGFRe5TpJ03CidMHMGyg6AwzU0G50Sblv5G07kRshzGoVHaBoC9poQAvD_BwE

## Slide 53: CellInsight CX7 High-Content Screening (HCS) Platform
https://www.thermofisher.com/us/en/home/life-science/cell-analysis/cellular-imaging/high-content-screening/high-content-screening-instruments.html

## Slide 55: Digital microscopy with automated brightfield and fluorescence imaging
https://www.moleculardevices.com

## Slide 56
http://zeiss-campus.magnet.fsu.edu/articles/livecellimaging/imagingsystems.html

## Slide 57
https://www.leica-microsystems.com

## Slide 58
https://www.nexcelom.com

## Slide 59
https://www.olympus-lifescience.com/en/solutions/live-cell-imaging/

## Slide 60
https://www.bruker.com/products/fluorescence-microscopes/fm-campaigns/live-cell-imaging.html

## Slide 61
http://www.mvi-inc.com/products/electronic-imaging/nikon-imaging-systems/

## Slide 62: TOMOCUBE INC
Hepatocyte in necrosis captured in 4D by the HT-1 holographic microscope
http://www.tomocube.com/product/ht-microscope/
Revolutionary holotomographic (3D holographic) microscopy opens new era for label-free live cell imaging

## Slide 63: Lecture Summary
Live cell applications are relatively common using confocal microscopy
Correct use of fluorescent probes necessary
Temperature and atmosphere control may be required
Thick specimens often require advanced image processing
Exotic applications are potentially useful
A limited window of time is available to image live cells before cells deteriorate