## Slide 1: BMS 631 - LECTURE 10 Flow Cytometry: Theory
Lynn Hall G221
Purdue University
Office: 494 0757
email; robinson at flowcyt.cyto.purdue.edu
WEB http://www.cyto.purdue.edu
Data Structure and Basic Multiparameter Data Analysis
J. Paul Robinson SVM Professor of Cytomics
Professor of Biomedical Engineering
Purdue University
Notice: The materials in this presentation are copyrighted materials. If you want to use any of these slides, you may do so if you credit each slide with the author’s name.
3rd Ed. Shapiro p 207-214
4th Ed p
This lecture last updated: March 23, 2016

## Slide 2: Data Analysis
File Structure
Gating
Data displays
  histogram
  dot plot
  isometric display
  contour plot
  chromatic (color) plots
  3 D projection

## Slide 3: Progression of Flow Cytometry
1 sample
per hour
At best
1960s
1970s
1980s
1990s
2000s
2010s
3-4 sample
per hour
5-10 samples
per hour
20-60 samples
per hour
50-100 samples
per hour
2000-20,000 samples/Hr
File
standards
2 color
1 color
3-8 color
5-18 color
Hyper spectral
Parallel μ-fluidics
Floppy
PDP 11
computer
100 to 500 megs
500 megs to
100 gigs
Tb or >
No
software
Manuf.
System software
3rd Pty
software
3rd Pty
software
3rd Pty
software
Lab automation
Data
pipelines
Notes: On the left is the only output of the original coulter counter – a graph on an oscilloscope – this was the data you had – eventually Fulwylers 1965 cell sorter (the original invention) has an osciloscope – and the only way to get the data off was with a polaroid camera – and one could actually cut out the histogram and weight it…!

## Slide 4: Data Acquisition
Each measurement from each detector is referred to as a “parameter” or “variable”
Data are acquired as a “list” of the values for each “parameter” (variable) for each “event” (cell)
[RFM]
Notes:     .Data Collection
Most flow cytometers collect between 3 and 10 "parameters" or "variables". Each variable can be used to discriminate some component of the cell populations. The standard for data collection has been known as the "Flow Cytometry Standard" or FCS data file structure. Data can be collected in either histogram mode or listmode. It is now routine to collect flow cytometry data in listmode because there are few constraints in the size of data collection files. Originally, flow cytometers did not have good network capability, small hard disks (5-20 Mb, and floppy drive that could save only about 1 Mb of data). With high capacity hard disks and Ethernet access now routine, listmode is the most useful collection modality. Data can be collected in a gated or ungated mode. If the data are ungated, all events registered by the instruments electronic circuits will be saved in listmode files. If gated collection is chosen, it is possible to collect a listmode file of only those events that satisfy certain criteria established by the operator.

## Slide 5: Flow Cytometer Computer Files
Listmode files
-correlated data file where each event is listed sequentially, parameter by parameter
-large file size
Histogram files
    uncorrelated data used for display only
Flow cytometry standard (FCS1.0 (1984) FCS 2.0 (1990), FCS 3.0 (1997), FCS 3.1 )2010)
    format used to save data
    use other software programs to analyze data
MIFlowCYt (2008)
Minimum Information about a Flow Cytometry Experiment
http://isac-net.org/Resources-for-Cytometrists/Data-Standards/Data-File-Standards/Flow-Cytometry-Data-File-Format-Standards.aspx

## Slide 6: Types of Listmode data
Ungated Listmode (normal for all instruments today) this means everything is collected
Gated Listmode collection
  (this is essentially no longer done because storage is not a big issue as it was when disc drives were only 5 megabytes)
Notes:     .Data structure
As noted above the file structure of the collected data is normally saved in a format known as FSC (Flow Cytometry Standard). This was originally proposed by Murphy et al (ref) to encourage all manufacturers and software engineers to comply with the standard to allow relatively easy data exchange between users of different instruments. Most manufacturers comply with a reasonable extent and a number of conversion utilities exist in software packages to allow cross-analysis of data from different instruments. The current standard is FCS-2, which adds a number of additional components to the file structure.
Each file collected has a number of keywords, PMT (Photomultiplier Tube) values and gating regions associated with it. Some listmode file structures also contain the instrument setup used to collect the original data. While each instrument collects files of different internal structure, in essence the information collected retains a relatively common format allowing 3rd party software manufacturers to create value added analysis packages.

## Slide 7: Data Acquisition - Listmode
Structure of a list mode file
Notes:     .List Mode Data
Essentially however, a listmode file is a series of numbers which sequentially list data for each variable collected on each cell. This is shown in the associated table provided.

## Slide 8: Listmode File Example
FILEVERSION;1.15
BTRIEVE;7058418 IBA13-2 collected 6/27/95;28/6/1995;;2;1
PARAM;LS1;400;LS2;600;FL1;460;FL2;400;FL3;300;WID;TIM
GAIN;1;1;2;2;0
THRES;0;10;0;5;0;5;1;11;0;5
FLUIDIC;3;4;1;7.20
FCM;357;2;2;1;1;0;0;1;1;1;1;0;1;1;0;1
SUBTRACT;15;0
AUTOCYCLE;1;0;0;0;0;0;0;-1;1;20000;7;0;1;0;1;0
PCBUFFER;374000;1
SERIAL;2
ADBOARD;1;0;0;30;0
PEAKAREA;0;0;0;0
CALIBRATION;300;300;300;300;0;0;0;0
MAINWINDOW;1
PRINTHEADER; Flow cytometry Report;Win-Bryte software - Bryte-HS flow cytometer
PRINTFOOTER; Purdue University Cytometry Laboratories
PRINTSTATISTICS;1;1;1;1;1;1;1;1;1;1;1;1;1;1;1;1;1;1;1;1;1;1;1;1;1;1;0;0;1;20;15;1;8;12;5;0;0;5
ROI;1;1;3;3;4;45;65;61;65;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;bird;1;1;1;0;1;1;1;0;1;1;0;;-1;-1;3;0;255;0;0
ROI;2;1;3;3;0;121;41;141;41;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;human;1;1;1;0;1;1;1;0;1;1;0;;-1;-1;3;0;255;0;0
ROI;7;3;5;1;0;0;18;24;63;52;48;16;2;0;2;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;0;;1;1;1;0;1;1;1;0;1;1;0;;-1;-1;3;5;255;0;0
ROIDATACYTO;7;20022;4845.7;98.7;0.0
ROIDATAHIST;1;2.4;16712;52;53;4044.6;2.4;53;83.5;0.0
ROIDATAHIST;2;2.2;1875;128;128;531.3;1.7;128;7.8;0.0
HISTOGRAM;FALS;24;49;290;320;256;0;Count;1;1;0;0;0;255;0;255;0;0;0;0;0;255;1;1;C;0;1;1;1;S;0;-1;-1;-1;-1;-1;-1;-1;-1;-1;-1;-1;-1;-1;1;1;1;255;0;255
HISTOGRAM;SIDE-SCATTER;290;49;556;320;256;1;Count;1;1;0;0;0;255;0;255;0;0;0;0;0;255;1;1;C;0;1;1;1;S;0;-1;-1;-1;-1;-1;-1;-1;-1;-1;-1;-1;-1;-1;1;1;1;255;0;255
CYTOGRAM;LS1-LS2;24;320;290;591;64;1;0;1;1;0;0;0;255;0;255;0;0;0;0;0;0;1;1;O;0;1;0;0;S;0;6;-1;-1;-1;-1;-1;-1;-1;-1;-1;6;-1;-1;1;0;1;0;A;B;C;D;1;1;255;0;0;32;32;0;255;0;255
HISTOGRAM;FL2;290;320;556;591;256;3;Count;1;1;0;0;0;255;0;255;0;0;0;255;0;0;1;1;C;10;1;1;1;S;0;0;1;-1;-1;-1;-1;-1;-1;-1;-1;0;7;-1;1;1;1;255;0;255
CYTOGRAM;FL2-TIM;556;320;822;591;64;6;3;1;1;0;0;0;255;0;255;0;0;0;0;0;0;1;1;I;0;1;0;0;S;0;-1;-1;-1;-1;-1;-1;-1;-1;-1;-1;-1;-1;-1;1;0;1;0;A;B;C;D;1;1;255;0;0;32;32;0;255;0;255
LISTDATA
226;0;426;426;243;81;0
412;225;8;426;426;239;0
380;262;2;427;427;245;0
578;348;0;412;412;239;0
529;377;0;431;431;240;0
420;203;0;438;438;243;0
444;221;0;427;0;240;0
383;215;0;421;421;238;0
735;716;0;1027;1027;280;0
499;228;0;431;431;239;0
531;328;0;433;433;241;0
520;218;0;423;423;243;0
471;252;0;425;425;244;0
652;298;0;441;441;240;0
561;291;0;421;421;240;0
608;307;0;415;415;241;0
541;231;0;424;424;245;0

## Slide 9: Listmode data Analysis
226;0;426;426;243;81;0
PARAM;LS1;400;LS2;600;FL1;460;FL2;400;FL3;300;WID;TIM
LS1 LS2 FL1 FL2 FL3 WID TIM
226 0 426 426 243 81 0

## Slide 10: Analog vs digital data collection
Data from BD White paper on Diva (Ben Verwer dated 2001)

## Slide 11: Digital structure (18 bits)
Data from BD White paper on Diva (Ben Verwer dated 2001)

## Slide 12
The famous “picket fence”
Data from BD White paper on Diva (Ben Verwer dated 2001)
9 values
90k values

## Slide 13: Gating
Real-time gating vs. software gating
Establishing regions
Gating strategies
Quadrant analysis
Complex or Boolean gates
Back gating
Note: some software differentiates between a region and a gate…
To me it is 6 of one thing and half a dozen of the other.....

## Slide 14: Real-Time vs. Software Gating
Real-time or live gating:
Restrict the data that will be accepted by a computer (some characteristic must be met before data is stored)
Software or analysis gating:
Excludes certain stored data from a particular analysis procedure

## Slide 15: Establishing Regions
Establishing regions:
  -objective or subjective?
  -training/skill/practice
Possible shapes:
  -rectangles
  -ellipses
  -free-hand
  -quadrants
Statistics
R1
Notes:     .Gating: Principles and Terms
There are a number of terms associated with flow cytometry analysis. One important concept is the idea of creating an electronic gate. A gate can also be referred to as a bitmap or region. Usually these terms are synonymous in usage, although they can mean slightly different things to different manufacturers.

## Slide 16
Region 1 established
Gated on Region 1
Using Gates
log PE
R1

## Slide 17: Quadrant Analysis
log PE
(+ +)
( - +)
(+ -)
(- -)

## Slide 18: Drawing Regions: Sample Preparation
B.subtilis spores B.subtilis veg. + spores
Debris
Spores
Spores
Debris
Vegetative
Data removed
From analysis

## Slide 19: Complex or Boolean Gating
With two overlapping regions, several options are available:

## Slide 20: Boolean Gating
Not Region 1:

## Slide 21: Boolean Gating
Not Region 2:

## Slide 22: Boolean Gating
Region 1 or Region 2:

## Slide 23: Boolean Gating
Region 1 and Region 2:

## Slide 24
Not (Region1 and Region 2):
Boolean Gating

## Slide 25
0
200
400
600
800
1000
Side Scatter Projection
Light Scatter Gating
Forward Scatter Projection
90 Degree Scatter
Neutrophils
Lymphocytes
Monocytes
Forward Scatter
Human white blood cells (red cells lysed)
☺
☺
☺
Projections show 2D
Plots in 1 parameter
space

## Slide 26: Back-Gating
Region 4 established
Back-gating using Region 4
log PE
Back gate

## Slide 27
3 Parameter Data Display
Isometric Display

## Slide 28: Methods that can change results:
1. Doublet discrimination
2. Time as a quality control parameter
Example: DNA content
- need to eliminate debris & clumps
- need to gate out doublets
- maintain constant flow rate

## Slide 29
DNA Histogram
G0-G1
S
G2-M
Fluorescence Intensity
# of Events
Time
Counts
A
B
C
Gating out bad data
Possible
blockage

## Slide 30: Time as a quality Control Parameter
A
B
“blockage” or other fluidic problem - data can be removed by gating the time histogram.
Abnormal histogram caused by turbulence (not biological variation)
Normalized histogram after subtraction of list mdoe data collected during the blockage shown at B on right.
A
B
Notes: Time was first proposed as a quality control parameter by Watson. IN the above example the DNA histogram (Top left) appears to have an abnormal G0-G1 population. By collecting time as a parameter, an observation can be made as to the rate of cell collection. As can be seen at the right of this histogram, the collection rate appears to have been significantly reduced at region B. This may have been caused by a blockage in the flow cell or some other non-biological problem. If this caused turbulence in the flow cell, a significant impact might be seen on the CV of the fluorescence histogram. By removing the data collected during the time gate B (by using a Boolean operator [a NOT B], only data collected at the appropriate rate is included in the data analysis as seen on the lower left histogram. A normal time histogram is shown on the lower right.

## Slide 31: Multi-color studies generate a lot of data
1 2 3 4 5 6 7 8 9 10
3 color
4 color
5 color
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+

## Slide 32
Contour plots
Dot plots

## Slide 33: Typical phenotypic analysis histograms
This figure shows two examples of simultaneous 2 color immunophenotyping. In figure 3 (a) the directly labeled MABs used were CD4-PE / CD8-FITC. In this example approximately 50% of the cells were positive for CD4 and 23% positive for CD8. These percentages were calculated based upon the settings of the negative control for 2% positivity. Right figure shows a similar situation for CD2-PE / CD19-FITC.

## Slide 34: Kinetic Analysis
50 ng PMA
Stimulated
Fluorescence
Fluorescence
0 ng PMA
Unstimulated
TIME (seconds)
0
1800
450
900
1350
TIME (seconds)
0
1800
450
900
1350
Figure: This figure shows an example of stimulation of neutrophils by PMA (50 nm/ml). On the left the unstimulated cells show no increase in DCF fluorescence . On the right, activated
cells increase the green DCF fluorescence at least 10 times the initial fluorescence.
References: (1) Robinson, J. Paul, Durack, Gary & Kelley, Stephen: "An innovation in flow cytometry data collection & analysis producing a correlated multiple sample analysis in a single file". Cytometry 12:82-90,1991.
(2) Durack, Gary, Lawler, Gretchen, Kelley, Stephen, Ragheb, Kathy, Robinson, J. Paul: "Time Interval Gating for analysis of cell function using flow cytometry" Cytometry, 12:701-706, 1991
References: (1) Robinson, J. Paul, Durack, Gary & Kelley, Stephen: "An innovation in flow cytometry data collection & analysis producing a correlated multiple sample analysis in a single file". Cytometry 12:82-90,1991.
(2) Durack, Gary, Lawler, Gretchen, Kelley, Stephen, Ragheb, Kathy, Robinson, J. Paul: "Time Interval Gating for analysis of cell function using flow cytometry" Cytometry, 12:701-706, 1991

## Slide 35: Color Coded Dot Plots
Key to understanding this figure, is the notion that different populations can be identified by colors and the relationship of these populations to one another can be monitored.

## Slide 36: Contour Plot with Projection
90 Light Scatter
60
50
40
30
20
10
0
Forward Scatter
60
50
40
30
20
10
0
Similar to the dot plot with projection
90 Degree Scatter
0
200
400
600
800
1000
Lymphocytes
Monocytes
Neutrophils
Side Scatter Projection
Forward Scatter
200
400
600
800
1000
8
15
20
30
40
50
100
200
1000
Scale

## Slide 37: 3 Color Combinations What does it mean?
Negatives
Positives
4+4=8
FITC
PE
APC
4+4+4=12
PE
FITC
FITC
PE

## Slide 38: 3 Color Combinations
FITC
PE
APC
4+4+4=12
PE-FITC Pattern
FITC
TR-FITC Pattern
FITC
TR-PE Pattern
Phycoerytherin
.1
1
10
100
1000
.1
1
10
100
1000
.1
1
10
100
1000
.1
1
10
100
1000
.1
1
10
100
1000
.1
1
10
100
1000
CD4
CD5
CD38
K/L
CD45
CD3
Negative
Negative
CD10
HLA-DR
CD20
CD8
CD7
CD2
CD8
CD8 (dim)
CD2
Negative
CD4
CD5
K
Texas Red
Texas Red
Phycoerytherin
CD3
CD3
IgM
IgG
CD3
CD8
CD3
CD20
IgG
L
CD20
CD20
CD45
These are actually 3D relationship, not 2D as may appear on this figure

## Slide 40: Lasers used for multicolor studies

## Slide 43: Many color flow cytometry
Nature Reviews Immunology 4, 648-655 (August 2004) | doi:10.1038/nri1416
Seventeen-colour flow cytometry: unravelling the immune system
Stephen P. Perfetto, Pratip K. Chattopadhyay & Mario Roederer

## Slide 44: Innovative Data Analysis

## Slide 45
The trick was the invention of a virtual parameter “Tube ID” Or Tube Identification Parameter (TIP) and a Single file structure for an entire assay
Reference: Robinson, J. Paul, Durack, Gary & Kelley, Stephen: "An innovation in flow cytometry data collection & analysis producing a correlated multiple sample analysis in a single file". Cytometry 12:82-90,1991.

## Slide 46
Dozens of files linked during collection
Robinson et al, Cytometry 12:82-90, 1991
14 Tubes, 2 color, single file, single plot

## Slide 47
Robinson et al, Cytometry 12:82-90, 1991
14 tubes, 3 color – single representation – expandable to n colors

## Slide 48: Innovation
Technologies Licensed to Beckman Coulter in 1991

## Slide 49: Use of time as a parameter for gating

## Slide 50: Creation of data sets for rapid analysis
132 samples
per hour
ANALYSIS
132 sample in about 2 minutes
Circa 1991
(reducing to DB file
This file structure was adopted by “Intellicyt” and is the basis of their combination files used for HT Flow (but was developed by us in 1991)

## Slide 51: Advanced ways of showing data relationships
20
60
100
Enrico Lugli et al, Università di Modena e Reggio Emilia
Oral Presentation Immunology section 15.30-17.30 today
Classification
?

## Slide 53: Spectral analysis allows classification

## Slide 54
Cytometry Pt A, Volume: 97, Issue: 10, Pages: 1044-1051, First published: 24 August 2020, DOI: (10.1002/cyto.a.24213)

## Slide 55
Cytometry Pt A, Volume: 97, Issue: 10, Pages: 1044-1051, First published: 24 August 2020, DOI: (10.1002/cyto.a.24213)

## Slide 56: Optical Design of a basic flow cytometer
PMT 1
PMT 2
PMT 5
PMT 4
Dichroic
Filters
Bandpass
Filters
Laser
Flow chamber
PMT 3
Scatter
Sensor
PMT 1
PMT 2
PMT 5
PMT 4
Dichroic
Filters
Bandpass
Filters
Laser
Flow chamber
PMT 3
Scatter
Sensor
Sample
HyperCyte diagram kindly supplied by Larry Sklar
HyperCyt system added to flow cytometer
Copyright 2006 Nature Publishing Group
Nature Reviews Protocols
384 wells/10 min
1 ml/sample
Notes: Basic schematic for how a flow cytometer works – a cell passes via a laser (blue line) and signals are passed down to each PMT and light scatter detector

## Slide 57
Historical References:
Robinson, J. Paul, Durack, Gary & Kelley, Stephen: "An innovation in flow cytometry data collection & analysis producing a correlated multiple sample analysis in a single file". Cytometry 12:82-90,1991.
Durack, Gary, Lawler, Gretchen, Kelley, Stephen, Ragheb, Kathy, Robinson, J. Paul: "Time Interval Gating for analysis of cell function using flow cytometry" Cytometry, 12:701-706, 1991
9 color very fast flow cytometer
High speed HyperCyte robot

## Slide 58
HyperPlex =
HyperCyt and Luminex
Theoretical potential for 50 plex in 1536 well format, 10 min
(20M per day per detector)
Selectivity
20 bead set,
2 colors, 7 levels each
Color 1, 7 levels
Color 2, 7 levels
Slide kindly supplied by Dr. Larry Sklar

## Slide 59
Automated Cytomat
Incubator
Automated Biomek
Preparative system
Automated HyperCyte
Sample Delivery
Cyan flow cytometer

## Slide 60: 384 Well Plate layout For drug screen – Mito Memb Pot
Valinomycin
positive controls
Negative
control wells
Negative
control wells
JC-1
2 color assay
3-5,000 cells per well (1 to 2 x 106 cells/plate)

## Slide 61: Plate Analyzer – all the parts

## Slide 62: Instantly see any well or combination of wells with any parameter or combination of parameters

## Slide 63
All 384 list mode files are loaded in a few seconds…

## Slide 64
Highlight the Plate Heatmap Format

## Slide 65
Highlight the Histogram/Dotplot Format

## Slide 66
Highlight the 3D Display Format

## Slide 67
Highlight the IC 50 Display Format

## Slide 68
Highlight the Analysis Logic Format

## Slide 69
Direct IC50 generation (on 4 functional tests simultaneously)
IC50-1
(32 drugs)
Conclusion from this:
The goal is to design assays as systems and integrate into the design an analytical process with a defined result. The focus of such approaches is a known endpoint.

## Slide 70
Data from every graph is available as an XL spreadsheet

## Slide 72
Combinatorial gate
Any Boolean gate
Gate any well or combination of wells
Interactive gate
On any global parameter
Any single parameter gate
Any 2 parameter gate
6 IC50 Models, Log/lin conversion
Possible Gate Strategies

## Slide 78: SPADE-Spanning tree Progression of Density normalized Events
(SPADE) is an algorithmic visualization tool for high-dimensional flow and mass cytometry data.
http://docs.flowjo.com/d2/plugins/spade/
http://blog.cytobank.org/2011/05/20/what-is-spade/
Bendall SC, Simonds EF, Qiu P, Amir ED, Krutzik PO, Finck R, Bruggner RV, Melamed R, Trejo A, Ornatsky OI, Balderas RS, Plevritis SK, Sachs K, Pe’er D, Tanner SD, Nolan GP. Single-cell mass cytometry of differential immune and drug responses across a human hematopoietic continuum. Science. 2011 May 6; 332(6030):687-96

## Slide 79: Some free software

## Slide 80: Learning objectives
Understand the data structure of flow cytometry data files
Understand the basic concepts for gating populations of cells
Understand the relationships between different parameters
Understand back-gating
Understand the basic principles of multiparameter analysis
Observe how High-Throughput flow (HT Flow) can be achieved

## Slide 81: Conclusions
The more parameters you have, the more complex the analysis will be
But…when you have more parameters (variables) you have more opportunities for population discrimination
Display of data in histogram and dotplot formats assists the analysis process
Displays in 3D are nice but not particularly useful for analysis.
Multiple parameter displays such as PCA or LDA are more useful for high content data sets
Very high throughput (96 to 384 well plates) require special treatment – new forms of data processing are necessary