A. Blech

1. It requires no more than a couple of minutes of bleach to remove the PI from the sample line of the
FACSCalibur.

2. An extensive but normal wash is fine. If in question, run some  stained cells through just to be sure the fluorescence is at background before running a patient or experiment through. It is Acridine Orange that can bind to tubing and cause problems, not PI.

3. 2-5 minutes 0.5% sodium hypochlorite (diluted "Household Bleach" in USA, there is something else in Europe that has chlorine used for cleaning?), 2-5 minutes 70% etanol and 2-5 minutes water

4. Use 50% bleach for 10minutes, followed by two coulter clenz wash each for 2 minutes and rinse with DIH2O for 4 minutes two times.

5. To wash with a solution of bleach (around 10%) for about 5 minutes followed by water for 5 minutes.

6. Run through with 10% bleach or undiluted one for one minute with SIP arm on the side, then center the arm and continue to run for another 5 minutes. Replace bleach with distilled water, run one minute with the SIP arm on the side, then another 10 minutes with the arm centered. It works well for us.

7. Bleach, 10% bleach. Soap will be nearly impossible to wash out completely. its residue will lyse cells in the next samples. Run 3 ml through at the end of each day, followed by 3 ml of distilled water. Let the SIP sit in dH2O when not in use.

8. Run at least 10mls of 5% bleach, followed by 5mls 5% Sodium Hydroxide and then 5 mls deionised water.

B. Others

1. When we were doing lots of DNA work with PI, we used 2%SDS to rinse instead
of water after the last sample. We often had to do the maintenance in SDS
instead of bleach, also.

2. Use 50% DMSO in water. Works great also for cleaning out proteins.