## Slide 1: Development Progress and update IP Portfolio Microparticle Detection Micro Imaging Hyper-spectrum detection
Jun. 7, 2016
M. Yamamoto
Keegan Hernandez
Paul Robinson
Confidential

## Slide 2: IP Portfolio Priority
Scanning Image Flow Cytometry
Top priority : PCT US/EP/JP
CAPTL only
Particle Detection
Using reflective surface
Flow Cytometry using Hydrodynamically Planar Flow
Time Sequential Flow Cytometry
2nd priority : PCT Joint or US only
Single Photon
Multi wavelength
Photo-reactive Cytometry
Hamamatsu Joint
PCT filed 12/19/15
US
- US filed 5/15/13
-US Issued 6/21/16
-JP filed 11/13/15 by 5/12/17
-EP filed 12/15/15 amnd by 6/15
- PCT filed 4/1/16
Phase Imaging Flow Cytometry
US
by 6/20/16
Scanning illumination format
Reflective surface format

## Slide 4: Optical Phase Detection for Nanoparticle Analysis
Workshop 11: Nanoparticle Characterization
June 13, 2016
M. Yamamoto
Paul Robinson
Notes: We would like to explain optical phase detection for nanoparticle analysis as new approach.

## Slide 5: Blue sky and Newton Ring
Scatter intensity:
- Essential principle of Flow Cytometry:
- relative value with calibration bead
- λ/2n : one way path (Transmission)
- λ/4n : two way path (Reflective)
- defined by wavelength
λ/4 = 100nm at 405nm
Notes: Blue sky scattering is gold theory for flow cytometry. On the other hand, light wave has another principle known well as “ Newton Ring”.
Newton ring is interference phenomena defined light wave phase difference. Reflective phase difference pitch is a quarter wavelength
equal to 100nm at 405nm.

## Slide 6: Blu-ray Read-Out
<Blu-ray Parameters>:
Reflective with 100um cover layer n=1.5
wavelength : 405nm NA= 0.85
Spot size : 250nm (FWHM)
Track pitch : 320nm control accuracy <±10nm
Pit depth : 60nm
Pit length : 150nm min.
A
B
C
D
HF = A+B+C+D
RPP =(A+B) – (C+D)
TPP =(A+D) – (B+C)
Blu-ray Standard Drive
Read Power = 0.30 mW
Track
Notes: Newton ring principle is widely applied for optical disc since CD, DVD and Blu-ray.
Your Blu-ray disc drive is reading and controlling several 10 nm accuracy with 250nm spot size and 0.3mW light intensity.
Reflected spot is detected by 4-splitting photo detector as HF, radial and tangential pushpull signal.
Based on Blu-ray development experience, we considered to apply Blu-ray platform for nanoparticle detection.

## Slide 7: Gold particle Read signal on 500nm pitch Land & Groove
Au 50nm HF & RPP
Au 50nm SEM picture
Au 30nm SEM picture
Au 30nm HF & RPP
Land
250nm
Groove
250nm
Land
250nm
Groove
250nm
Depth=40nm
Spot Size FWHM
250nm
High spatial resolution is confirmed
Notes: Preparing 250nm width and 40nm depth Land and Groove substrate, gold particle is tested.
Shown on picture, signal shows clear pulse with good S/N even for 50nm or 30nm.
Gold particle is optically black body, so you can understand why detectable by 250nm spot.

## Slide 8: Micro channel with 2um height for liquid sample
Sample Inlet
Substrate 1.1mm
Cover 100um
Channel Gap 2um
Microchannel zone
with reflective grating
Label
100um cover
Sample inle
0.5mm dia.
Microchannel
width 2mm
Depth 2um
Label side
Read Out side
Notes: Next step was transparent particle detection in 3 dimensional liquid medium.
On the reflective Land and Groove substrate, the ring shape channel is manufactured with 2um height.
Sample inlet is from 1.1mm substrate with 0.5mm diameter.
Microchannel is covered with 100um cover layer.

## Slide 9: Particle Read out signal in Liquid(Water)
Gold particle with large k modulates HF,RPP and TPP.
Transparent particle is detected by TPP signal.
Signal level is proportional to Δn x d.
Gold 100nm signal in water
Polymer 100nm signal in water
500nm
Notes: Picture shows gold and polymer 100nm particle signal in water.
Transparent particle does not modulate HF signal, mainly detected as TPP signal.
Traverse signal pitch is equal to 500nm and valley to peak correspond to 250nm.
From simulation based on Vector Theory, signal amplitude is proportional to refractive index difference and particle size.
Thanks to high signal to noise ratio, it is possible to analyze each pulse amplitude, width and phase.

## Slide 10: Firstly observed plasma sample signal: diluted x10 in water
Channel radius=20.0 -22.0mm
Area size = 264 mm2
Channel Height= 2 um
Channel Volume= 0.5 ul
If concentration = 109 /ml
Estimated MV in channel= 0.5 million
HF
TPP
Detected waveform
Reference: A Flow Cytometric Method for Characterization of Circulating Cell-Derived Microparticles in Plasma
Notes: Recently, we tried to evaluate human plasma sample processed by centrifuge described on reference paper.
Channel radius and estimated concentration suggests to include half million particles in microchannel.
Picture shows variety of pulses and confirmed OPD with a quarter micron spot has very high sensitivity and spatial resolution.

## Slide 11: Actual disc read
Signals on Oscilloscope
Pioneer have provided us with drive control access under an NDA
Actual disc reading of particles on our custom made blu-ray discs
Off shelf Pioneer Blu-ray disc with our electronic modification board
Below is a read of an actual disc with particles (see trace above)

## Slide 12: Liquid Biopsy Work Flow
Sample IN
Whole blood
Cell Separation
staining
Microvesicles
Measurement
Sorting
Liquid biopsy
Slide glass shape LOC
Disk shape LOC
R-θ scanning
X-Y scanning
PCR
Targeted
Sequencing

## Slide 13: 8 sample disk layout by modified DVD replication
Vacuum channel ?
Measurement zone
Sample Inlet
Any number of channels can be created
In any shape

## Slide 14: Summary & Steps
Optical phase detection looks powerful for nanoparticle analysis.
Established mass production can provide low cost platform.
3. further steps;
Identification of pulse signal origin
Fluorescence photon detection from antibody
Signal analysis algorithm and statistics
Embedded scale and Calibration method
Disc quality improvement and production
Drive modification and optimization
Format development for nanoparticles
Simulation by Vector Theory FDTD
Others for diagnostics platform establishment
Notes: On the other hand, there are a lot of issues to be resolved.
The most important issue is how to identify each particle.
In order to identify, antibody and fluorescence photon detection is mandatory.
Fortunately, we developed single photon detection technology and possible to integrate with Blu-ray optical pick up.
We believe reflective phase detection with smaller spot is promising alternative to analyze nanoparticles.
Thanks for your attention and interest.

## Slide 15: Time Domain Analysis by Scanning Small Laser Beam Cytometry
Parallel 15: Flow Cytometry Instrumentation
June 14, 2016
M. Yamamoto
Paul Robinson
Notes: Thank you for kind introduction.
Today’s my talk is technical challenges for flow and imaging integration and single photon detection for next generation cytometry.

## Slide 16: Micro Imaging Flow System Concept :
P-signal
Laser
Photon
Detector
Deflector
Ob. Lens1
Ob. Lens2
S-signal
Fl signal
B.S.
Signal
Processor
Memory
CPU
Flow Chamber
Chart: VCO-IN[Vpp]
| 0.0 | 0.1 | 0.2 | 0.3 | 0.4 | 0.5 | 0.6 | 0.7 | 0.8 | 0.9 | 1.0 | 1.1 | 1.2
 | 0 | 4.1 | 7.9 | 11.6 | 15 | 18.5 | 22 | 25.2 | 29 | 32 | 35.9 | 39.4 | 43.1
 | 0 | 3.9 | 7.4 | 11 | 14.2 | 17.5 | 21 | 24.3 | 27.6 | 30.5 | 33.7 | 37.8 | 40.6
 | 0 | 4 | 7 | 10.5 | 13.8 | 16.8 | 20.4 | 23.6 | 26.6 | 29.9 | 33.1 | 36.5 | 39.3
 | 0 | 3.9 | 7 | 10.1 | 13.3 | 16.4 | 19.7 | 23.1 | 26.2 | 29.4 | 32.3 | 35.5 | 38.7
100kHz
500kHz
750kHz
1000kHz
1.80μmFWHM
What useful output
for cellular analysis?
Sheath flow
Sheath flow
Laser Scanning
Hydrodynamic
Focused Plane*
(Confocal view)
Sample
16bit/160MHz/8CH ADC
20Gbps/3TB
Wavelength stabilized 405nm laser
Notes: In order to achieve small spot size and required speed, several new devices are developed
Wavelength stabilized 405nm laser with min. wave front aberration, acoustic optical deflector and
Full digital driver, custom objective lens and planar flow chamber. Small spot and planar flow is possible to achieve high throughput even at 1m/s flow velocity.
It is possible to detect many signal from cell - transmission, polarization, diffraction and fluorescence.
Data acquisition system has developed for the highest performance.

## Slide 17: Dimensional Analysis at Scan Direction:
Laser spot
with scan velocity v
Object
Distance= Δt x v
Detected Signal
Saw tooth
Scan drive
Scan direction
Flow direction
10um scale image
Laser illumination with scale
Developed Digital Driver
Notes: Regarding to scan direction dimensional analysis, length is basically measured by spot scanning velocity and pulse width. Spot position accuracy is calibrated with microscope calibration pattern with 10um.
Previous VCO was used analog mixer, but spot accuracy was not sufficient, then full digital driver is developed.
As the result, carrier level and frequency accuracy is much improved. In addition, it is possible to add scale marker in illumination beam.

## Slide 18: Dimensional Analysis at Flow direction :
Channel Wall
Center Spot
Side Spot -1
Side Spot +1
Distance
d
d
Flow velocity
v
Parabolic Velocity
3 Spot aligned
in parallel to
flow direction
Laser
Deflection
3-split detector
Phase Grating
f
Det(-1)
Det(0)
Det(+1)
Flow
Scanning 3 spot by grating
Spot separation d = 35um
Notes: Regarding to flow direction, it is possible to calculate by averaged flow velocity.
But, flow speed in micro channel is in general parabolic shape and sometimes complicated status depend on channel design. More accurately, 3 spots generated by phase grating are useful.
Time difference for each particle can provide microscopic velocity and precise particle dimension at flow direction.

## Slide 19: Time Sequential Observation of Target Cell by 3 Spot and Slide Motor
Scanning Optics
Sample IN
Δｔ
2Δｔ
3Δｔ
nΔｔ
Δｔ=0
flow velocity=v 0.1 ~ 1.0 m/s
Slide displacement by linear slide
Sheath
Sheath
Detection Optics
Particle flows parallel to channel wall
“Cell Tracking”
Notes: Combination of 3 spot optics and high speed linear motor can provide new feature to evaluate specific cell in time sequentially. We named as Cell Tracking. Cell tacking is useful to monitor rapid chemical and physical change of cell under chemical or physical exposure.
Picture and movie shows experimental set-up.

## Slide 20: Fluorescence Single Photon Detection :
Ex. 405nm Photon/sec = 3.06 eV = 0.489E-18 [W] ~ 0.5aW
633nm Photon/sec = 1.95 eV = 0.313E-18[W] ~ 0.3aW
780nm Photon/sec = 1.59 eV =0.254E-18{W] ~ 0.25aW
(Hamamatsu Handbook)
Notes: Photon detection is very important for flow cytometry. In usual, we use MESF refer to standard molecules or Q&B for instlement evaluation. Once back to SI unit and Einstein/Planck equation, single photon has specific energy depend on wavelength.
According to simple calculation, visible wavelength photon energy is 3 to 1.5 eV. This means 2 or 4 phtons per second has 1 aW energy.
aW ( E-18) is not usual unit, but please remind 2 blue-violet photon per second has 1 aW energy as order understanding.

## Slide 21: Photon Detector(1) : microPMT by Si MEMS Process
High Speed
TIA Preamp
~100MHz
Photon Pulse Width 4 ~ 5ns
Small photocathode 3x1mm
Photocathode material
Si process accuracy
High Speed &
Low dark count
Chart: Dark Count at HV=-950V
| 0.0 | 5.0 | 10.0 | 15.0 | 20.0 | 25.0 | 30.0 | 35.0 | 40.0 | 45.0 | 50.0 | 55.0 | 60.0 | 65.0 | 70.0 | 75.0 | 80.0 | 85.0 | 90.0 | 95.0 | 100.0 | 105.0 | 110.0 | 115.0 | 120.0 | 125.0 | 130.0 | 135.0 | 140.0 | 145.0 | 150.0 | 155.0 | 160.0 | 165.0 | 170.0 | 175.0 | 180.0 | 185.0 | 190.0 | 195.0 | 200.0
 | 82.412 | 13.318 | 1.152 | 0.746 | 0.458 | 0.384 | 0.198 | 0.174 | 0.106 | 0.06 | 0.054 | 0.018 | 0.018 | 0.018 | 0.008 | 0.02 | 0.002 | 0.004 | 0.008 | 0.004 | 0.002 | 0.006 | 0 | 0 | 0 | 0 | 0.004 | 0.002 | 0 | 0 | 0 | 0 | 0.002 | 0 | 0.008 | 0 | 0.006 | 0 | 0 | 0 | 0
[mV]
[count/sec]
Notes: PMT is the key detection device for flow cytometry. Recently silicon MEMS technology is applied for microPMT by Hamamatsu. Micro-PMT has much smaller photocathode and dynode area. Thanks to smaller size effect, microPMT shows very low dark count at room temperature and high speed photon pulse response.
There are device deviation, but one of our devices shows several count per second. This means it is possible to evaluate aW level light intensity without cooling.

## Slide 22: Photon Detector(2) : Si Photo Multiplier by Geiger mode
Very High Speed
TIA Preamp
100MHz ~ 1GHz
Chart: Count/sec
| 0.0 | 5.0 | 10.0 | 15.0 | 20.0 | 25.0 | 30.0 | 35.0 | 40.0 | 45.0 | 50.0 | 55.0 | 60.0 | 65.0 | 70.0 | 75.0 | 80.0 | 85.0 | 90.0 | 95.0 | 100.0 | 105.0 | 110.0 | 115.0
10s Gate | 528 | 3717 | 87493 | 90787 | 93212 | 95155 | 95962 | 92961 | 1758 | 1335 | 1268 | 1244 | 1228 | 1218 | 1120 | 923 | 51 | 23 | 22 | 20 | 20 | 18 | 11
40s Gate | 21 | 20 | 19 | 16 | 11 | 3 | 1
Pulse width 600~900ps
Dark count & PE characteristics
0.5PE
1.5PE
2.5PE
Notes: Another advanced photon detector is Si photomultiplier with same internal gain to vacuum PMT.
Combined with ultra high speed preamplifier, it is possible to observe shorter than 1ns photon pulse.
SiPM is still higher dark count, but it is possible to analyze picoseconds time difference.
SiPM is now applying photon counting medical imaging to detect small tumor.
We can also apply these photon detectors to flow cytometry.

## Slide 23: “Photon dots” and Bead Phase Image on slide
Bead phase image (405nm scan)
Single photon dots
as background noise
40um
10um marker
Bead
Photon pulse
Notes: On the other hand, we immediately met tough hurdle to overcome.
This image is scanning phase image of blank bead and single photon map detected by microPMT.
Each white dot shows background phton caused by auto fluorescence in optics.
In order to analyze accurately, the biggest issue is to achieve AFL free optics.

## Slide 24: Background Analysis(1) : Laser Background Photon Spectrum
Chart: 405nm laser spectrum
| 352.29 | 363.08 | 373.86 | 384.65 | 395.44 | 406.23 | 417.02 | 427.81 | 438.6 | 449.39 | 460.18 | 470.97 | 481.75 | 492.54 | 503.33 | 514.12 | 524.91 | 535.7 | 546.49 | 557.28 | 568.07 | 578.8599999999999 | 589.64 | 600.4299999999998 | 611.22 | 622.01 | 632.8 | 643.59 | 654.38 | 665.17 | 675.9599999999999 | 686.74 | 697.53 | 708.3199999999999 | 719.11 | 729.9 | 740.69 | 751.48 | 762.27 | 773.0599999999999 | 783.8499999999999 | 794.63

 | 6010 | 7550 | 3330 | 5790 | 61600 | 202000 | 52000000 | 1110000 | 127000 | 225000 | 1080000 | 7250000 | 342000 | 74700 | 18400 | 8900 | 5790 | 1800 | 3470 | 2540 | 4470 | 3620 | 2440 | 718 | 440 | 4710 | 367 | 483 | 5250 | 735 | 3130 | 2710 | 2240 | 1160 | 5670 | 980 | 3880 | 4310 | 2070 | 1860 | 1030 | 3440
Laser Peak
E+11
10pW
10fW
10aW
Origin: ASE, AFL
Notes: One of background photon sources is laser itself. We are usually concerning to laser wavelength, but laser beam includes low level but wider spectrum photon by ASE or AFL..
Level is E-6 or less to laser wavelength, but need to control for single photon analysis.
As shown in figure, photon counting method can evaluate extremely low level noise from pW to aW level.

## Slide 25: Background Analysis(2) : Objective Lens objective AFL illuminated by 405nm/ 440LP OD12 : F5.6/1”
Reflecting Objective
20x FL objective lens
MPlan 20x
Chart: Photon Count/s
| 352.29 | 363.08 | 373.86 | 384.65 | 395.44 | 406.23 | 417.02 | 427.81 | 438.6 | 449.39 | 460.18 | 470.97 | 481.75 | 492.54 | 503.33 | 514.12 | 524.91 | 535.7 | 546.49 | 557.28 | 568.07 | 578.8599999999999 | 589.64 | 600.4299999999998 | 611.22 | 622.01 | 632.8 | 643.59 | 654.38 | 665.17 | 675.9599999999999 | 686.74 | 697.53 | 708.3199999999999 | 719.11 | 729.9 | 740.69 | 751.48 | 762.27 | 773.0599999999999 | 783.8499999999999 | 794.63
 | 766.6667 | 815 | 451.6667 | 433.3333 | 358.3333 | 325 | 356.6667 | 603.3333 | 390 | 29200 | 20100 | 16633.3333 | 14400 | 17833.3333 | 20100 | 12500 | 15800 | 16500 | 16200 | 22800 | 25000 | 32700 | 24000 | 48700 | 35600 | 15700 | 11300 | 17600 | 19500 | 18166.6667 | 6516.6667 | 7700 | 3050 | 5283.3333 | 1625 | 1616.6667 | 1175 | 528.3333 | 1046.6667 | 715 | 1138.3333 | 1253.3333
~20fW
Notes: Another background photon source is objective lens. This picture shows objective output illuminated by 405nm.
Photon spectrum is shown as histogram with peak level 20fW.
Objective AFL is caused by variety of optical glass with different index and coating.
Objective lens is inevitable for fluorescence optics, but also noise source illuminated by shorter wavelength.
Laser noise is possible to eliminate with clean up filter, but AFL free objective is the challenging including optical materials.

## Slide 26: Background Analysis(3): Optical Coating and Substrate
Al 10nm +HC (R~50%)
Al 50nm +HC (R~90%)
Al 50nm only ( low level Trans)
Al 100nm only
Borosilicate Glass +dust
Float Soda Glass +dust
ISO6400/F6.3/1.0sec
Notes: 3rd background source is sample substrate and optical coating. Even very thin AR or hard coating, it has relatively large AFL.
Slide glass material is Soda lime or borosilicate specified low AFL, but it emits strong AFL from single photon perspective. Metal reflective layer has much lower than dielectric coating.

## Slide 27: Excitation Ratio=Detected Photon/Illuminated Photon
Coating | PE Count A*exp(-kx) | Excitation Ratio/ 2.04E+14 |  | (405nm/100uW)
Al 10nm +HC | 6,750 | 3.30E-11 | include glass AFL | 
Al 50nm + HC | 2,340 | 1.14E-11 | caused by HC | 
Al 50nm only | 510 | 2.50E-12 |  | 
Al 100nm only | 340 | 1.67E-12 |  | 
Quartz | 590 | 2.89E-12 |  | 
“NanoBlack” | 200 | 0.98E-12 |  | Lowest level
Bz. Card | 5.74E+7 | 2.81E-7 |  | Highest level Almost saturated
Highlighter | 1.79E+7 | 1.14E-7 |  | 
Notes: In order to evaluate AFL for variety of material, excitation ratio is defined as detected photon versus illuminating photon.
Due to limitation of collected solid angle, total AFL level may be x10 larger.
As known, highlighter or fluorescence material has E-6 order of excitation ratio.
Optical material and coating has E-11 to E-12 level with clearly shown difference.
Currently min. excitation ratio is 1E-12 – may be limited by long pass filter characteristics.

## Slide 28: Newly developed “NanoBlack” by CNT :
Polymer surface: 4% Reflectivity “NanoBlack” < 0.1% & AFL free
(Courtesy by Purdue Nanotech Center)
Vantablack (UK) is similar CNT surface with reflectivity < 0.035%.
Notes: We are now looking for AFL-free material in UV ~visible region.
One of interesting material is carbon nanotube developed by Purdue nanotech center.
You can look much lower reflectivity than plastic surface. Vantablack is similar black-hole surface as the world darkest. Carbon was not considered as optical material, but now we need to check again from AFL- free perspective.

## Slide 29
Chart: Light Power & Number of Photon (405/488/561/638nm)
| 0.001 | 0.0001 | 1.0E-5 | 1.0E-6 | 1.0E-6 | 1.0E-7 | 1.0E-8 | 1.0E-9 | 1.0E-10 | 1.0E-11 | 1.0E-12 | 1.0E-12 | 1.0E-13 | 1.0E-14 | 1.0E-15 | 1.0E-16 | 1.0E-17 | 1.0E-18

 | 2037150000000000 | 203715000000000 | 20371500000000 | 2037150000000 | 2037150000000 | 203715000000 | 20371500000 | 2037150000 | 203715000 | 20371500 | 2037150 | 2037150 | 203715 | 20371.5 | 2037.15 | 203.715 | 20.3715 | 2.0372

 | 2454640000000000 | 245464000000000 | 24546400000000 | 2454640000000 | 2454640000000 | 245464000000 | 24546400000 | 2454640000 | 245464000 | 24546400 | 2454640 | 2454640 | 245464 | 24546.4 | 2454.64 | 245.464 | 24.5464 | 2.4546

 | 2821830000000000 | 282183000000000 | 28218300000000 | 2821830000000 | 2821830000000 | 282183000000 | 28218300000 | 2821830000 | 282183000 | 28218300 | 2821830 | 2821830 | 282183 | 28218.3 | 2821.83 | 282.183 | 28.2183 | 2.8218

 | 3209140000000000 | 320914000000000 | 32091400000000 | 3209140000000 | 3209140000000 | 320914000000 | 32091400000 | 3209140000 | 320914000 | 32091400 | 3209140 | 3209140 | 320914 | 32091.4 | 3209.14 | 320.914 | 32.0914 | 3.2091
mW
uW
nW
pW
fW
aW
kilo
Mega ( 1 photon/us)
Giga
Tera
Peta
Conventional Fl level
Excitation
Photon
E+16~E+17
“ Photon Flow Chart”
change unit to photon number
Background Photon,
“ From Noise to Frontier ”
Notes: This is simple graph between light power and number of photon based Planck equation.
From 10mW illumination to detectable limit of dark count is now 16 or 17 decades.
Conventional flow cytometry detects fluorescence from pW to nW range. 1pW includes a million photons.
This is reasonable because of one background photon per microsecond.
Advancement of detection now enable expanding noise detection limit to atto watt level.
We believe previous noise region is now changing to frontier of cytometry for micropartcle detection, non-labeling, single molecular analysis.

## Slide 30: 42CH Hyper-spectrum Detection

## Slide 31: 42CH Experimental Set-Up

## Slide 32: Miftek Road Map
Shift forwarding to Business Phase
Define Market and Product
Development /Productization Issues for Product Specification
Collaboration & Funding Scheme/Time Line/ Business Plan
Start Project

## Slide 33: Business Category/Field
Module Business
Instrument Business
System(Hard-Soft-Disposable)
Business
Format & Clinical Business
Photon Detection module/software
Data acquisition module
Hyper-spectrum Flow Cytometer
Micro imaging Cytometer
MV analysis platform
(BD,BC,Philips,Sony etc.)
Clinical platform
(Roche,Biomerieux,etc.)

## Slide 34: Required MV Platform Resources:
Roche
Drive
Media
substrate
Material
Clinical
Reagent
Filter
Preprocessing
Control Software
Media Format
Analysis Software
etc.