## Slide 1: Flow Cytometry applications for antimicrobial resistance
J. Paul Robinson
Distinguished Professor of Cytometry
Dept of Biomedical Sciences
Professor of Biomedical Engineering
Purdue AMR conference Feb 27, 2024
wombat@purdue.edu

## Slide 2: Outline of the presentation
Overview of bacterial detection methods
Why flow cytometry for microbiology?
Bacterial analysis using conventional flow cytometry
Challenges in bacterial analysis using flow cytometry
Difference between conventional and spectral flow cytometry
Bigfoot spectral cell sorter and ELS technology for detecting MRSA
Measuring bacterial autofluorescence for detecting MRSA using bigfoot
Safety precautions
Opportunities for collaboration

## Slide 3: Bacterial detection and identification methods
Notes: Limitations of current method for bacterial detection

## Slide 4: General principle of flow cytometry in microbiology
Blood Culture
Petri Dish
3-4 Hours
Add dye
Run samples (~1 min each)
Flow Cytometer
Result

## Slide 5: Basic flow cytometry analysis
Antibiotic Sensitive
Antibiotic Resistant
10.3889/oamjms.2017.104
2 hours to result
2017 Study
How long did this take?

## Slide 6
“The time required to obtain a final AST, using FASTvet panel was 2H”
https://www.frontiersin.org/journals/microbiology/articles/10.3389/fmicb.2020.01944/full

## Slide 7
https://www.nature.com/articles/s41598-021-82186-4
Streptococcus pneumoniae
0.125 μg/mL penicillin
1 μg/mL cefotaxime
2 μg/mL cefotaxime
Media
SYTO9 and propidium iodide PI
PSSP, penicillin susceptible S. pneumoniae
PRSP, penicillin resistant S. pneumoniae
MRSP, multi-drug resistant S. pneumoniae
R
MR
MR
Similar data with Haemophilus

## Slide 8
https://www.thelancet.com/journals/ebiom/article/PIIS2352-3964(22)00326-7/fulltext
“Flow cytometry-assisted Antimicrobial Susceptibility Test (FAST) returns results in 3-6 hours.”
2022 Study

## Slide 9: Bacterial analysis using conventional flow cytometry
5 minutes

## Slide 10: Challenges in bacterial analysis using flow cytometry
Small size/volume of bacteria- signal to noise ratio.
Lack of availability of specific antibodies (stain) to all bacteria species/across strains.
Effect on stains on bacterial viability for downstream analysis.
Wide emission band among stains- signal overlap.
Autofluorescence from bacteria and media.
Need to safe sort pathogenic bacteria for clinical applications.
Safety
Antibodies
Volume of lymphocytes: 206 µm3 [1]
Volume of bacteria: 0.4-3µm3 [2]
Image credit: https://blenderartists.org/t/10-000-x-magnification/693576
Size/volume
References:
[1] Kuse R, et al, Blood lymphocyte volumes and diameters in patients with chronic lymphocytic leukemia and normal controls. Blut. 1985 Apr50(4):243-8
[2] Levin PA, Angert ER. Small but Mighty: Cell Size and Bacteria. Cold Spring Harb Perspect Biol. 2015 Jun 8;7(7):a019216.

## Slide 11: Working principle of cell sorting
Single cell handling capabilities
High throughput screening
Efficiency sorting
Multiplexing detection capabilities
Can be integrated to different down stream analysis
Available in clinical settings
Working principle
Why flow cytometer?
Notes: Why flow cytometry is interesting?

## Slide 12: Motivation
Can we address some of the limitations?
Wide emission and signal overlap.
Autofluorescence from bacteria and media.
Need for safe sorting pathogenic bacteria and identification.
Safety
Pathogens
?
Spectral signatures can address signal overlap.
Multiplexing capabilities
Sorter with biosafety features can be used for safe handling of pathogens.

## Slide 13: Spectral over conventional flow cytometers
Spectral flow cytometer [2]
Light emission across the whole spectra
More information with unique spectral signature
Stain 1
Stain 2
Conventional flow cytometer [1]
Specific band of emission wavelength recorded-
Loss of information due to signal overlap and spillover
[1] ThermoFisher, Fluorescence Spectral Viewer
[2] Images adopted from hand-book ‘’Particle Separation Techniques: Fundamentals, Instrumentation, and Selected Applications’’ (Robinson, 2022)
530/30 band pass filter

## Slide 14: Thermo Fisher’s Bigfoot spectral cell sorter features
Biosafety cabinet
Voltage plates
3D printed agar plate holder
Placed in normal laboratory settings
1.
3.
1. Anti-splash patrol
Photo credit: Thermo Fisher
2.
2. Contaminated air patrol
3. Sample exposure patrol
Three levels of border control:

## Slide 15
6 samples handling
9 lasers and 54 detectors
Bigfoot spectral cell sorter features continued..

## Slide 16: Detection and identification of antibiotic resistant bacteria using Bigfoot cell sorter and ELS technology

## Slide 18: Rapid single bacterial cell patterning using Bigfoot
Screenshots of sorting parameters in Bigfoot
96 well plate map
(52 spots)
Sorting and patterning polymicrobial population
Message: Sorting and patterning time depends on bacterial concentration and flow rate
Real time patterning of bacteria on agar plate
13 secs to plate 52 single bacterial cell (<1000 cells per sec)

## Slide 19: Preparation of Antibiotic Gradient Agar (AGA) plate
Bryson, V. and W. Szybalski, Microbial Selection. Science, 1952. 116(3003): p. 45-51
Layer1 :TSA agar Oxacillin
Layer 2: TSA agar
High
Low
3mm thick mat
Notes: Method for patterning antibiotic gradient plate (AGA) plate. Preparation of AGA plate in the hood by preparing two separate layers

## Slide 20: Patterning bacteria on agar plate using Bigfoot
Escherichia coli
Klebsiella pneumoniae
Staphylococcus aureus
Escherichia coli
(51 colonies)
Salmonella Enteritidis
(49 colonies)
Staphylococcus aureus
(52 colonies)
Klebsiella pneumoniae
(49 colonies)
52 colonies/plate
96 well plate
(reference)
Mapped wells/spots
After 24hr
Notes: Here the idea is to show single cell patterning on the agar plate using Bigfoot sorting using 96 well plate maps. 52 colonies per plate.
Then we show live dead (heat killed) bacteria patterned on agar plate

## Slide 21
Methicillin-susceptible staphylococcus aureus
(MSSA)
TSA agar
0ug/ml
Low
High
AGA agar
Total 1ug/mL
AGA agar
Total 2ug/mL
Layer1 :TSA agar Oxacillin
Layer 2: TSA agar
High
Low
Oxacillin concentration
Incubation time: 15hr
Incubation time: 11hr
Patterning antibiotic susceptible & resistant bacteria on antibiotic gradient plate
Methicillin-resistant staphylococcus aureus
(MRSA)
Notes: Examples of patterning MSSA and MRSA on Oxacillin AGA plate using Bigfoot.
Bacteria grown on TSA used as positive control
MSSA found growing at low conc region on 1 ug/mL and no growth on 2 ug/mL AGA plate after 15hr. (11 hr is enough to find this though!!)
MRSA found growing all over 1ug/mL AGA plate where as only at low conc region on 2ug/mL after 15hr

## Slide 22
AGA agar
Total 1ug/mL
Incubation time: 31hr
AGA agar
Total 2ug/mL
AGA agar
Total 4ug/mL
Low
High
Oxacillin Concentration
MRSA
MSSA
Patterning MSSA and MRSA on AGA plate (continued)
Notes: The plates from previous slide result were incubated to observe for colony formation
After 31hr MSSA growth found only at low to medium conc region on 1ug/mL AGA plate and low conc region on 2ug/mL AGA plate
After 31hr MRSA growth found everywhere on both 1 and 2ug/mL AGA plate and low to medium conc on 4ug/mL AGA plate

## Slide 23: Bacterial identification using Elastic light scattering
Tang Y et al., 2014

## Slide 24: Scatter patterns of different bacterial species
E.coli
(non-pathogenic)
TSA agar
SE
(non-pathogenic)
SE 015
(Pathogenic)
XLT4 agar
VRBG agar
2X magnification
TSA: Tryptic Soya agar
(Both gram negatives and positives)
Selective agars:
XLT4: Xylose-Lysine-Tergitol 4 agar
(Gram negatives, mostly Salmonella sp. and some E.coli strains)
VRBG: Violet red bile glucose agar
(Gram negatives)
11hr of incubation
Notes: Scatter patterns of diff bacteria grown on different types of agar

## Slide 25
Colony =1
MSSA
TSA-11hr (Positive control)
Colony =19
Colony =26
Colony =1
Colony =19
Colony =3
MRSA
MRSA
Colony =5
Colony =2
Colony =7
Total Oxa.= 1ug/mL
Colony =2
Colony =3
Total Oxa.= 2ug/mL
MSSA and MRSA colony scatter pattern comparison
High
MSSA
AGA-11hr
Colony =1
Colony =3
Colony =4
Total Oxa.= 1ug/mL
Less
Oxa.
conc
Notes: Scatter pattern from ELS of selected colonies of MSSA and MRSA on AGA plate (1 and 2ug/mL )
Scatter pattern of MSSA and MRSA grown on normal TSA plate used as positive control and both colony pattern looks the same
Scatter patterns of both MSSA and MRSA has changed in the presence of antibiotics (depends on the antibiotic conc). The growth is mainly along the vertical direction rather than horizontal to avoid antibiotics stress on agar plate
Need machine learning to better differentiate between the colony scatter patterns

## Slide 26: Machine learning model for bacterial classification
PCA 2
Training set
PCA 1
E.coli
SE
MSSA and MRSA
MRSA on 1 ug/ml AGA
SE
E.coli and SE
MSSA and MRSA
Test sample
(MRSA on 1ug/mL AGA)
Test sample
Model prediction accuracy for training set:
97% for E.coli
98% for SE
76% for MSSA
72% for MRSA
Model prediction accuracy for test set:
96% for MRSA colonies on 1 μg/mL oxa AGA
Notes: Machine learning software to train the knowing scatter patterns of bacteria.
Examples shown using PCA after training then with E,coli SE MSSA and MRSA.
The CV matrix (not shown here) gave good training accuracy for all bacteria (>70%).
Similar looking scatter patterns cluster together such as for E.coli and SE; MSSA and MRSA
When MRSA grown on 1u1g/mL AGA plate was given as a test set, the model predicted with 96% accuracy that the colony scatter pattern belongs to MRSA.
This shows thar our model working accuracy is great and I is possible to identify bacteria using scatter patterns

## Slide 27: Detection of antibiotic resistant bacteria using bacterial autofluorescence signature

## Slide 28
Method: Bacterial autofluorescence analysis
Advantages:
If unique signatures are obtained for different bacteria, bacterial identification can be performed without staining.
Saves money and time!

## Slide 29: Autofluorescence spectral signatures of stressed vs non-stressed bacteria
Full spectrum – 9 lasers 54 detectors
Selected wavelengths – 3 lasers 29 detectors
1
2
3
1
2
3
Intensity
Ex-Em wavelength (nm)
E.coli
Bacteria in TSB
Bacteria in 0.125ug/mL genta
Bacteria in 0.25ug/mL genta
Bacteria in 0.5ug/mL genta
(Non-stressed)
(Stressed with Gentamicin)
Ex-Em wavelength (nm)
FCS express 7.0
Notes: Comparing autofluorescence signature of stressed (diff concentration of gentamicin) and non-stressed E.coli. Unmixed data using TSB + antibiotics used as negative controls
The full spectrum (9 lasers and 54 detectors) from bigfoot is shown. But it is the average signature of 20,000 events per case. Meaning for each stressed and non-stressed case the average signature of autofluorescence is taken from 20,000 bacterial events and shown as a full spectrum here.
Among 9 lasers and 54 detectors, selected ex (349, 405 and 488-532 nm) and em wavelengths shows increase in autofluorescence for stressed E.coli (green- 0.5ug/mL genta). This corresponds to 3 lasers and 29 detectors.

## Slide 30: Gentamicin stressed vs non-stressed bacteria
PC1
PC2
PC1
PC2
Dimensionality reduction method - Principal Component Analysis (PCA)
Salmonella Enteritidis
E.coli
Increase in autofluorescence observed at selected wavelengths for bacteria stressed with high conc of antibiotics
Bacteria in TSB
Bacteria in 0.125ug/mL genta
Bacteria in 0.25ug/mL genta
Bacteria in 0.5ug/mL genta
(Non-stressed)
(Stressed with Gentamicin)
Cytospec 11.4
Notes: Autofluorescence signatures projected in the form of PCA plots of the stressed and non-stressed E.coli and Salmonella shown
As E.coli is stressed with high conc of gentamicin, a clear shift in the PCA density plot is seen showing increase in autofluorescence for the stressed E.coli and Salmonella

## Slide 31: Detection of MRSA by stressing with Oxacillin- PCA
MSSA
MRSA
2 hr
6 hr
Increase in autofluorescence observed at 6 hr for non-resistant strain.
Non-resistant or resistant strain can be detected in 6 hr.
PC1
PC2
Bacteria in TSB
Bacteria in 1 ug/mL oxa
Bacteria in 2 ug/mL oxa
(Non-stressed)
(Stressed with Oxacillin)
Notes: Detecting MSSA and MRSA by stressing them with oxacillin and measuring autofluorescence at selected wavelengths (3 lasers 29 detectors)
Autofluorescence signatures projected in the form of PCA plots, showing two time points (2 hr and 6 hr) comparing stressed and non-stressed MSSA and MRSA)
Bacteria gown in TSB used as positive control and TSB + antibiotics used as negative controls
At 2hr no major difference between stressed and the non-stressed case were observed for both MSSA and MRSA
At 6hr stressed MSSA (green -2ug/mL oxacillin) shows shift in the PCA density plot where as MRSA which is resistant shows no significant difference from growing in TSB media
By adding 1 or 2ug/mL Oxacillin, presence or absence of MRSA can be said by measuring autofluorescence from the suspected samples

## Slide 32: Why does autofluorescence increase in stressed bacteria?
Photo credit:
Nealson, K. H et al., (2016)
Increase in fluorescence is due to increase in flavin protein/FAD, NADH cofactors
(This mechanism is enhanced when exposed to stress such as antibiotics. Amount varies based on bacterial species)
Shewanellla oneidensis MR‐1
Bacterial metabolism
Increase in flavins,
FAD and NADH
Notes: Reasons for increase in autofluorescence.
When stressed they take Flavin pathway which results in increase in concentration of flavins, FAD and NADH leading to increase in autofluorescence as a part of survival.

## Slide 33: Safety precautions
Placed in a dedicated room for sample preparation
BSL2 cabinet
Goggles for splash protection
Lab coat
Double gloves
Dedicated box for transporting plates
Separate incubator for pathogenic organisms
Trained professionals only!
Normal lab settings
Trained professionals maintain the same dress code
Lab bench cover trashed after experiment

## Slide 34: Safety precautions - Contamination management
Waste tank closed cabin
After experiment we clean the Bigfoot system by running:
10% Bleach (5min)
70% ethanol (5min)
Coulter clenz (5min)
Cavicide solution (5min)
Check with milliQ water for contamination (10min)
Bigfoot runs a thorough cleaning step during shutdown
Separate waste tank for pathogenic bacteria
(Bleach, Overnight)
Non-pathogenic bacterial samples plated after cleaning with 70% ethanol – 48hr incubation.
Mouse and keyboard
Inside the bigfoot biosafety hood
Message:
Few months were spent to optimize the cleaning steps to prevent any sort of contamination before working with pathogenic bacteria.

## Slide 35: Summary and Conclusions
Flow cytometry is designed to run small particles and cells – bacteria area well suited
Because we can measure a very small change in organism number in a sample, we can determine if an organism is growing or not
With a specialty instrument like the Bigfoot sorter – we can sort highly pathogenic organisms
We can also sort and culture directly on to petri dishes or tubes
We can measure changes in autofluorescence which has potential for rapid identification of resistant organisms
BMS in the vet school has the most advanced flow cytometer currently available and we are interested in collaborating with anyone interested

## Slide 36: Acknowledgements
  Purdue team:
  Dr. Bartek Rajwa
  Dr. Euiwon Bae
  Dr. Sungho Shin
  Dr. Xi Wu
  Dr. Sharath Iyengar
  Kathy Ragheb
  Brianna Dowden
  Valery Patsekin
  John González
  Rémi Lechardoy
  Sanjana Shemonti
  Hyun Jung Min
  Joohyeon Lee
  Dr. Kennedy Okeyo
  Dal-Seong Yoon
  Noah Bousier
  Collaborators:
  Dr. Moon Kim
  Dr. James Lindsay
  Dr. Andrew Gehring
  Dr. Martha Vaughn
  Dr. Stephen On
  Dr. Jake Jacobberger
  Assoc Prof Bill Eades
  Dr. Andreas Cossarizza
  Dr. Euiwon Bae
  Dr. Bartek Rajwa
  Miftek Team
  Masanobu Yamamamoto
  Keegan Hernandez
  Daniel McMillan
  Zilin Xu
  Andy Bieberich
  Josh Harriger
  Jared Thompson
  John Gonzales
  Nikolas Mahaffey
  David McMillan
  Aamodini Jayaram
  Valery Patsekin
  Brianna Corman
  Kathy Ragheb
This research was supported by the Center for Food Safety Engineering at Purdue University, funded by the U.S. Department of Agriculture, Agricultural Research Service, under Agreement No. 59-8072-1-002 and Agreement No. 58-8042-0-061 to Purdue University Cytometry Laboratories. Any opinions, findings, conclusions, or recommendations expressed in this publication are those of the author(s) and do not necessarily reflect the view of the U.S. Department of Agriculture.
Corporate Support Partners:
ThermoFisher, Beckman Coulter
SciAps,
@Cytometryman
College of Veterinary Medicine