## Slide 1: Advanced sensitivity and resolution in flow cytometry through innovation
J. Paul Robinson
The SVM Professor of Cytomics &
Professor of Biomedical Engineering
Purdue University
Home to:
Purdue University Cytometry Laboratories
Cytometry Listserve: cytometry@lists.purdue.edu
Cytometryman

## Slide 2: What is different about this small machine?
Millipore/Guava
Apogee
Luminex
Accuri
Bryte
BC
BD
BD
BC
BC
BC
Amnis
Bay
Biosciences
BD Influx
BD Aria

## Slide 3: 4 Essential Feature differences
The detectors are different
The light management technology is different
The fluidics are different
Multiple side scatter options

## Slide 4: Why do we care about these things?
Detectors:
Because 100% of other cytometers use PMTs
PMTS have some advantages that have made them very successful in fluorescence detection for 40 years
Normally you stay with systems that work!

## Slide 5: Why do we care about these things?
Light Manipulation
Traditional dichroics work very well
Traditional bandpass filters also work well
Fiber optics also work well

## Slide 6: Why do we care about these things?
Fluidics
Positive pressure systems are excellent
Only downside maybe not quantitative
No major negatives

## Slide 7: Why do we care about these things?
Side Scatter
Traditionally off a blue laser on most systems
Works well for typical sized particles and cells
Need to consider how to look at smaller particles

## Slide 8: Cytoflex changes a few things…
Detectors – non standard
Optics – non standard
Fluidics – non standard
Mixed laser combinations

## Slide 9: Detectors

## Slide 10: Detectors
No PMTS – only APDs
PMT
APD

## Slide 11: APDs
The big deal about APDs is their incredible quantum efficiency around 80% compared to 25% for PMTs
the number of incident photons
the number of electrons emitted
Remember what QE is: =
Notes: Remember what QE is: it's the number of electrons emitted divided by the number of incident photons

## Slide 12: Avalanche Photo Diodes
Notes: Another big difference between an APD and a PMT is the bias voltage – you remember that PMTs typically require around 1000v whereas APDs are about 20 or 90 volts and it’s not continuous like the HV has to be on a PMT Absorption of incident photons creates electron-hole pairs. A strong internal electric field is created by a high reverse bias voltage This field accelerates the electrons through the silicon crystal lattice to
produce secondary electrons by impact ionization. The resulting electron avalanche produces a high gain.

## Slide 14: Wavelengths

## Slide 15: Detection optics

## Slide 16: Optical Fibers

## Slide 17: Wavelength Division Multiplexing
Modern communication systems can handle up to 160 signals and can thus expand a basic 10Gbit/s system over a single fiber pair to over 1.6 Tbit/s
Transponder
Transponder

## Slide 18: WDM

## Slide 19: How does WDM Work?
Input light
From fiber optic
Band pass filters
Mirror
To detector (APD)
420/15
450/15
480/15
520/15

## Slide 20: WDM

## Slide 21: WDM Splitter

## Slide 22: Double WMD
3
3

## Slide 23: Simplicity, efficiency
Highest efficiency of all current optical pathways for flow cytometers
Its requires very small physical space so is compact
Absolutely minimal distance from signal to detector
Requires a tiny bandpass filter
Any bandpass filter can essentially be put into any of the WDM modules
Virtual unlimited range of combinations of lasers and output pathways.

## Slide 24: Fluidics

## Slide 25: Peristaltic Pump

## Slide 26: Fluidics
Time
Fluorescence

## Slide 27: James Watson “Time” as a parameter on EVERY sample collected
Cytometry 8:648-649, 1987
Fluorescence
Time
Fluorescence

## Slide 28: V-SSC Side Scatter

## Slide 29: Light Scatter
What do we know?
We know that generally the larger the difference in the refractive indices, the more light will be scattered by a particle

## Slide 30: Light Scatter
We also know that the more granular the particle is, the probability of getting light reflecting off of subparticle components increases.
That is why we always see granulocytes have such high side scatter compared to say lymphocytes that have no granules.
We also know that the amount of light scattered by any particle is proportional to the diameter of the particle and inversely proportional to the wavelength of the light being used to detect it.
FSC
SSC
Lymphocytes
Monocytes
Granulocytes
RBCs, Debris,
Dead Cells
FSC
SSC

## Slide 31: Light Scatter
So if we lower the wavelength of the excitation, we will see increased side scattered light scattering at any given particle size – if we move from say 488 nm excitation to 405 nm excitation, this will increase the range of resolution of smaller particles than can be detected by standard side scatter.
We also know that when entering a medium of a different refractive index, light waves are refracted by the new medium inversely proportional to the wavelength of the light, with lower wavelengths having a higher refraction than larger wavelengths
Thanks Isaac Newton!!!

## Slide 32: Violet Side Scatter (VSSC)

## Slide 33: Multiple Side Scatter detectors
CytoFLEX supports 2 side scatter (90 deg) channels. One from 488nm laser, which is designed for cells (minimum detection size 200nm) and the other from violet laser, designed to detect particles as small as 80nm.

## Slide 34: Lasers

## Slide 35: CytoFLEX S with Near UV Laser
Laser portfolio for the CytoFLEX S
  Near UV (375 nm)
  Violet (405 nm)
  Blue (488 nm)
  Yellow/Green (561nm)
  Red (638 nm)

## Slide 36: Lasers
405 nm
488 nm
561 nm
638 nm

## Slide 37: WDM for 4 lasers including NUV
405 nm
561 nm
375 nm
488 nm

## Slide 38: WDM for 4 lasers including NUV
405 nm
375 nm
638 nm
488 nm

## Slide 39: CytoFLEX S Near UV Laser
375 nm Near UV laser
  Power: 60 mW
  Spot size: <5µm x 80µm
CytoFLEX S portfolio with the near UV laser

## Slide 40: System layout with the Yellow Laser

## Slide 41: Why use a 375 nm laser?
375 nm laser is as good or BETTER for the excitation of BV dyes, commonly excited using the 405 nm laser.
375 nm laser allows for applications such as DAPI live/dead and Hoechst 33258 used in murine side population research.
375 nm is an excellent alternative to expensive “true UV”, 351 nm laser.
375 nm does not produce the high thermal output found in the UV 351 nm laser from the laser head or power supply.
375 nm laser is as good or better for the excitation of BUV dyes, commonly excited using the 355 nm laser

## Slide 42: Let’s look at BUV dyes with 375 nm Laser
Whole blood stained with CD3 + DAPI
BUV 737 with a 780 or 712 nm bp filter

## Slide 43: 12 color experiment including BUV

## Slide 44: Laser Combinations

## Slide 45: System Resolution

## Slide 46: Rainbow Beads Analysis

## Slide 47: Advantage of 561 nm laser

## Slide 48: Viability with Hoechst and DAPI
1 day old blood stained with DAPI or Hoechst

## Slide 49: Side Population courtesy of Dr. Bill Telford

## Slide 50: What might the big Take home messages be….???
Because of the incredible linearity of the APD detectors:
Can use a CD4-stain in a comp-tube and apply the results to any multi-color assay containing the same stain on a dim antibody
The compensation elements of any stable chromophore, e.g., non-tandems such as FITC, PE etc., can also be re-used forever
A user can store these in a compensation library, then mix and match them with other freshly prepared elements for unstable chromophores to construct a full comp matrix.

## Slide 51: The Holy Grail of flow…..
A next generation of detectors will have self-calibration capability such that users will be able to port calibration matrix from one instrument to another, across the entire CytoFLEX platform.

## Slide 52
Questions