# SOP-E003

Shut-Down Procedure for FC500-Gallios Flow Cytometer

Objective:  To clean and shut down the FC500-Gallios.

Procedure:

- Prepare 5 tubes as follows:

  - bleach

  - clenz

  - Water

  - Water

  - Water

- You will have to run each one individually, manually, so put tube on just like a sample

- Run each of the above tubes for about three minutes each

- On the FC500 software, put the machine in standby mode by clicking on the ZZZZZZ button.

- Fill each of the two vacuum line cleaning tubes with water

- Load them onto the carousel in postion 1 and 2, with the “bucket” facing toward the center of the carousel.

- Then select the button which looks like the vaccum line cleaning tubes

- When it is finished sucking up the water from the two tubes, it will go into cleanse mode.

- When it is finished and says “press idle to continue”, just exit out of the software

- On the FC500 desktop, click on the “FC OFF” icon. The instrument will shut off

- On the Gallios software, exit out and close all windows.

- Turn off the power strip.

-----------------------------------------------------------------------------------------------------------------------------

Created by: Kathy Ragheb

Verified by:			 	Date:

Print Name

Sign Name