# SOP-E004

#### Adjustment of Flow Rate Slider Bar on the Beckman-Coulter CyAn ADP

Objective: To obtain optimum flow rate with enough “play” left in the slider bar movement. May need to do this after re-installing the software.

Procedure:

- Get some beads going which are at the typical concentration for doing QC.

- Adjust the slider bar to where you get about 150 beads/s.

- Remember what the pre-set number is.

- Go to Edit > preferences > instrument > CyAn

- You will see the boxes for minimum and maximum pre-sets.

- Type in the minimum pre-set you were supposed remember.

- For the maximum pre-set, enter 10,000 over the minimum.

- Commit and Save the changes.

--------------------------------------------------------------------------------------------------------------------------------

Created by: Kathy Ragheb						Date:  February 22, 2013

Verified by:			 	Date:

# Text boxes

Sign Name

Print Name