# SOP-E008

MoFlo XDP Sorter Daily Shutdown Procedures

Purpose:  to properly shutdown the XDP sorter at the end of the day

Procedure:

In general, power to the XDP system and aXcess Control Panel should remain on

- Turn off lasers

  - Turn off lasers at the remote controller boxes

  - Turn off laser power supplies, keys and switches

  - Turn off main power switches ( silver boxes on wall with red handles)

  - Allow laser cooling water to run about 10 minutes longer to cool lasers

- After 10 minutes, turn off chiller water pump switches on wall next to sorter

- Turn off valves behind sorter

- Turn off chilled water supply valves in micro room (133D)

- On the Stream Configuration Screen, turn off the Drop Drive, Deflection Plates, and Stream Illumination.

- Place a tube of 10% bleach on the SmartSampler and run it for 3-5 minutes.

- Place a tube of Coulter Clenz on the SmartSampler and run for 3-5minutes

- Place a tube of DI water on the SmartSampler for 5minutes

- Place another tube of DI water on

- Press the Stop Fluidics button (on SmartSampler red button shaped like a droplet), to shutdown fluidics

- Close the red (vacuum)and blue (pressure) switches located on the pressure console, to stop pressure and vacuum

- Turn off the aXcess control panel monitor ( button on the side of the monitor)

- Turn off Jun-Air compressor (main power switch on front, red)

- Loosen the vent valve on the sheath tank to de-pressurize the system NOTE: turn only one whole turn, otherwise you will liberate the compressed spring and ball bearing………good luck finding them!!!!)

-------------------------------------------------------------------------------------------------------------------------------

Created by: Kathy Ragheb,   August 2, 2011

Atoned by:

Verified by:			 	Date:

Print Name

Sign Name