# SOP-E011

MoFlo XDP Total Power (Hard) Shutdown and Restart

Purpose:  to properly shutdown the XDP sorter in the event of planned prolonged power outages, or when communications seem to be not working properly

Procedure:

- Follow the standard daily shutdown procedures [SOP-P073]

- Exit Summit software.

- Log off from windows and shutdown the computer, make sure the computer’s power is off

- On the Cytometer Control Panel also called the aXcess Control Panel, press the Controlled (via software) Shutdown Button (blue ) located in the lower right hand corner on the FRONT of the monitor (on the actual computer screen, not on the monitor frame).

- Turn off the XDP chassis power, black switch located midway on the back, on the right . (Next to the two yellow and two blue lights).

- Turn off the power to the other “black box” “online UPS” (Un-interupted Power Supply)

- NOTE: ONLY TURN THIS OFF FOR PROLONGED POWER OUTAGE. DO NOT TURN OFF IF SIMPLY REBOOTING THE SYSTEM FOR COMMUNICATION ISSUES, WHEN POWER WILL BE RESTORED IMMEDIATLEY.

To Restart after a “hard” Shutdown:

- Turn on the computer, and login

- Turn on the UPS unit , if it is off.

- Turn on the server

--------------------------------------------------------------------------------------------------------------------------------------------------------------------

Created by: Kathy Ragheb,   July 24, 2012

Atoned by:

Verified by:			 	Date:

Print Name

Sign Name