# SOP-E017

Upgraded MoFlo XDP Sorter Daily Shutdown Procedure

Purpose:  to properly shutdown the XDP sorter at the end of the day

Procedure:

In general, power to the XDP system and aXcess Control Panel should remain on

- Turn off lasers: 488 and 561 (underneath the optical bench)

  - Turn off the keys

  - Turn off power switches located on the back in the middle at the top of the power supplies.

- Turn off lasers: UV

  - Turn off the key, then the power switch located on the front of the power supply.

- Turn off lasers: 640 and 405

  - Turn off the power switches located on the cables on the optical table.

- If in sort mode: On the Stream Configuration Screen, turn off the Drop Drive, Deflection Plates(HV), and Stream Illumination.

- Place a tube of 10% bleach on the SmartSampler and run it for 3-5 minutes.

- Place a tube of Coulter Clenz on the SmartSampler and run for 3-5minutes.

- Place a tube of DI water on the SmartSampler for 5minutes.

- Place another tube of DI water on the SmartSampler.

- Press the Stop Fluidics button (on SmartSampler red button shaped like a droplet), to shutdown fluidics.

- Close the red (vacuum)and blue (pressure) switches located on the pressure console, to stop pressure and vacuum.

- Turn off the aXcess control panel monitor ( button on the side of the monitor) .

- Turn off Jun-Air compressor (main power switch on front, red).

- Turn off the vacuum pump by unplugging the cables (with the green tape).

- Loosen the vent valve on the sheath tank to de-pressurize the system (NOTE: turn only one whole turn, otherwise you will liberate the compressed spring and ball bearing………good luck finding them!!!!)

-------------------------------------------------------------------------------------------------------------------------------

Created by: Kathy Ragheb,   September 9, 2015

Atoned by:

Verified by:			 	Date:

Print Name

Sign Name