# SOP-E018

#### Laser Alignment – Upgraded MoFlo XDP Cell Sorter

Objective: To align the lasers on the XDP for optimal performance and measurements.

Procedure: (perform in the following order: 488, FS, violet, red, UV, yellow)

- Check the stream alignment first:

  - On the axcess control panel, press the Coarse Alignment button (located on the top left of the screen).

  - Press the Pinhole Illumination button to see the image of the pinholes (the small light bulb at the lower left of the camera field of view).

  - Using the top micrometer, turn slightly in a clockwise direction to bring the nozzle tip into view. Inspect it for dried saline. If needed, take a piece of lens paper and wrap it around a cotton applicator stick. Moisten with water and wipe the tip.

  - While observing the stream, turn the micrometer in a counter clockwise direction until you feel it give (approximately 62 turns).

  - You are looking for the center stream borders to be the same thickness all the way from the nozzle down.

  - If not you need to tweek it

- Start laser alignment with the 488 laser. You may have to go back to the stream again, then back to the 488 laser.

- Align the FS.

- For the remaining lasers, turn on the top one first and align, then bring the second laser in.

- For (405)violet/red(640):

  - Align violet first, keeping the red laser shuttered.

  - Open the red laser shutter, it should be right there, if not you will have to adjust the alignment screws of the laser itself.

- For UV/yellow(561):

  - Align UV first, keeping the yellow laser shuttered.

  - Open up the yellow laser, if not right on, then use the alignment screws on the yellow laser to adjust it in.

Pinhole 1 = 488

Pinhole 2 = 405/640

Pinhole 3 = UV/561

-------------------------------------------------------------------------------------------------------------------------------

Created by: Kathy Ragheb	DATE:	9/11/15

Verified by:			 	Date:

Print Name

Sign Name