# SOP-E019

CytoFlex Start-Up Procedure

Purpose:  for getting the CytoFlex ready to run samples

Procedure:

- Check sheath and waste tanks.

  - Fill sheath tank with CytoFLEX sheath fluid if needed

  - Empty waste tank and rinse, if needed

- Close the bleed valve on the top of the sheath tank.

- Turn on Jun-Air compressor (main power switch on front, red).

- Turn the red (vacuum)and blue (pressure) switches located on the pressure console, to the on position.

- Turn on the small vacuum pump by connecting the two power cables (have green tape on them).

- Turn on the axcess control panel (press button at the side).

- On the sorter touch screen, SmartSampler section, press the button (should be green) shaped like a droplet to turn on the sorter fluidics.

- Once system pressure reaches 60psi, turn on de-bubble and let de-bubble for a while. (Go and get some coffee or something.)

- Open the valve on back of optical table for three seconds, to bleed off bubbles in the lines.

- Turn on lasers: 488 and 561 (underneath optical bench)

  - Turn on main power supply switches on top middle back of power supplies.

  - Wait for warm-up to finish. It will say standby. Turn keys to ON position.

- Turn on laser: UV (underneath optical bench)

  - Turn on main power switch located on the front of the power supply.

  - Let warm up for at least 30 minutes before turning key to enable. (Yellow lights will blink until stable).

- Turn on lasers: 640 and 405

  - Turn on the power switches located on the cables on the optical bench.

- Run QC beads.

- Set up for sorting (see SOP P084).

------------------------------------------------------------------------------------------------------------------------------------------------

Created by: Kathy Ragheb,   September 8, 2015

Atoned by:

Verified by:			 	Date:

Print Name

Sign Name