# SOP-E021

CyAn/HyperCyt/Biomek

Objective: To make the HyperCyt sipper, the CyAn flow cytometer, and the Biomek robot all communicate in order to run a lot of plates .

Note: 	IP address of CyAn is:    128.210.61.239

IP address of HyperCyt is:   128.210.61.81

To find the IP address of the CyAn open a DOS wondow and type:     ipconfig

CyAn- the CyAn needs to be turned on in a certain order anytime the computer is turned off, or if there is a strange problem.

- Turn everything off.

  - The main power switch on the back of the CyAn

  - The main power switch on the right side of the SMS unit (sheath/waste tank system)

- Shutdown the CyAn computer and turn power switch off.

- Power up everything in this order (order is critical).

  - SMS power switch (right side)

  - CyAn main power switch (back of machine)

  - CyAn computer

HyperCyt- if the HyperCyt gets disconnected form the CyAn, the following steps need to be reset:

- Open HyperCyt Controller software

- Click on tools tab

- select HyperCyt Options

- Under Flow Cytometer Integration, need to check the box that says “enable flow cytometer integration”. It should be on Cytometer Cyan, if not select Cyan from the dropdown menu.

- save

-----------------------------------------------------------------------------------------------------------------------------

Created by: Kathy Ragheb

Verified by:			 	Date: 	04/04/2011

# Text boxes

Sign Name

Print Name