# SOP-E027

#### (Shortlist) Operation Procedure For Coulter Counter Model Z2

Objective: To determine the concentration of blood cells in an acqueous suspension.

Procedure:

1. Turn on the power switch, located on top, left, front of unit.

2. Lower the stage and remove the coulter clenz vial.

3. Place a counting vial containing millipore water onto the stage and rinse 5-10 times by moving the stage gently up and down.

4. Now replace the vial with a counting vial containing PBS from the dispenser.

5. On the data terminal press <Functions>

6. Move the cursor keys to <Prime Aperture>

7. Press <Start>, then press <Start> again.

8. The system will now take several minutes to prime.

9. When it is finished, the data terminal message screen will read <aperture primed>

10. Put a coulter counting vial containing fresh PBS on the stage and run a count for the background. It should be < 250.

10. For your sample, you must enter the size you wish to count.

11. On the data terminal press <Set-Up> until screen S1: SETUP –ENTER SIZE DATA is displayed.

12. Check that the aperture tube displayed is the one being used.

13. Select <Units> should  be set on (m.

14. Move to Set Upper Size Tu, enter the numerical value of the size below which you want to count

15. Move to Set Lower Size Tl, enter the numerical value above which you would like to count.

16. The ratio of upper to lower size should not exceed 3:1.

17. Count mode should be set to <between>, all particles equal to or exceeding the lower size setting, but do not exceed the upper size setting are counted

18. Press <Start> to initiate an analysis

19. If count was selected, screen A4 displays the count result.

20. Record the count.

21. With the 100m aperture installed, a count of 1950, would be read as 1.950 x 106/ml.

22. With the 50m aperture installed, you need to multiply the count by 5, so 1950 x 5 = 9750= 9.75 x 106/ml

22. If you wish to repeat the count, just press <Start> again

23. You should take three counts and average them.

Note: The 100m aperture counts the number of cells in 500l

The 50m aperture counts the number of cells in 100l.

---------------------------------------------------------------------------------------------------------------------------------

Created by: Kathy Ragheb	Date:  Septeber 15, 2004

Verified by:			 	Date:

# Text boxes

Sign Name

Print Name