# SOP-E040

#### Operation of pH Meter- VWR 9100

Objective: To instruct proper usage of the VWR 9100 pH meter.

NOTE: Never wipe the electrode bulb, and only use water to rinse the electrode.

Procedure:

- Turn on meter. The button is on the bottom row 2nd from the left.

- Must calibrate, by first removing the old calibration values.

  - Hit CAL then CLEAR (CLR blinks in window)

  - Hit CLEAR again, will see three bars

- Now remove electrode from the storage buffer and rinse with water.

- Insert the electrode into pH4.0 standard and press the READ button.

- Calibration for pH 4.0 standard is complete when the AUTO eye stops flashing.

- When 4.00 appears, remove the electrode from the pH 4.0 standard and rinse with water.

- Now insert the electrode into pH 7.0 standard and press READ.

- Calibration for pH 7.0 standard is complete when AUTO eye stops flashing.

- When 7.00 appears, remove the electrode from the pH 7.0 standard and rinse with water.

- Now insert the electrode into pH 10.01 and press READ.

- Calibration for pH 10.01 standard is complete when AUTO eye stops flashing.

- When 10.01 appears, remove the electrode from the pH 10.01 standard and rinse with water.

- Calibration is now complete.

- Press EXIT.

- Place electrode in your sample.

- Press READ.

- When the reading is stable, record the pH.

- Rinse electrode with water and place back into the electrode storage solution.

---------------------------------------------------------------------------------------------------------------------------------

Created by: Kathy Ragheb	Date:

Verified by:			 	Date:

Print Name

Sign Name