# SOP-E043

#### Filter Scans

Objective: Scan filters to check or confirm transmission spectrum, wavelength, and filter type of optical filters

Procedure:

- Log into Novell Network

- Make sure that USB2000 is plugged in (via USB) and connected to filter scanning box (via fiber)

- Open Spectra Suite icon on Desktop

- Make sure “Electric Dark Correction” box is unchecked

- Remove black unit cover

- Turn lamp on- Switch located on back of blue “LS-1” unit

- On computer, lower “integration time” until spectrum is not over-saturated (no flat line at top of peak).

- *Should be below 15msec, if not may be necessary to realign fiber (thick blue wire leading to black transmission detection unit) until not over saturated at 15 msec.

- Increase Boxcar width to ~5.

- Change Scans Average to 1,000.

- Turn lamp off

- Replace black unit cover.

  - Be sure not to hit blue fiber at any point before, during, or after scans.  It fits easily in a groove on the black unit cover.  Hitting the wire will misalign the unit.

- Wait for spectrum to flat line

- Click on black light bulb icon

- Remove back unit cover and turn light on.

- Replace cover

- Wait for spectrum to return to max value

- Click the yellow light bulb

- Click the blue “T” for transmission mode

- Open Cover

- Find and remove correct size filter holder for filter scanning.  It may be lying near unit.

- Using pen suction cup, gently place filter in filter holder.

- *Note the direction of arrow on filter side.  This is the direction the light must pass through the filter. Arrow should point toward blue fiber wire.

- Carefully secure filter in holder with ring nut- Use micro spanner wrench to tighten gently.

- *Do not over tighten or filter may break!

- Place optical filter holder back into post holder on the stage.

- Turn black stage rotation knob until rotating stage reaches 0º.

- Turn optical filter holder in the stage until reflected light on “LS-1” light output opening is centered.

- Secure optical filter holder in stage post holder by turning knob on side of stage until tight.

- Replace black cover over unit.  Be careful not to touch blue fiber wire

- Minimize Spectra Suite software

- Double click on Excel document “Blank Filter Transmission Sheet” on desktop

- Click “File” → Click “Save As”

- Rename file with filter information: Wavelength, Filter Type, 90 º scan and Serial # or Part # (if available) – Files are usually saved under T:\Filter Scans

- Return to Spectra Suite

- Click “Copy spectral data to clipboard” icon. (Looks like two sheets of paper.)

- Return to Excel document

- Click in first empty cell in the wavelength column and paste data from clipboard. (Right click→”paste” or type “ctrl” + “v”)

- Add filter information in “Filter Transmission Scan” section. May not have all information so fill in what is available.  Most information may be found on filter carrying case or bag.

- Save sheet information: “File” → “Save”

- Print “% Transmission sheet” and “Optical Density” sheet

- To do a 45º scan of the same filter, remove black unit cover and turn stage rotation knob to 45 º.

- Repeat steps from #26 on.  Be sure to note it is a 45º scan in the Excel file name and in the cell indicating filter angle

- Carefully remove ring nut with micro spanner wrench. Lay optical filter holder on table for stability and filter safety.

- Carefully remove filter with suction cup pen and replace in carrying case or bag.

- If more filters to scan, repeat from step #19 with new filter.

- If no more filters to scan, turn lamp off, replace cover, close Spectra Suite and Excel, and log off.

---------------------------------------------------------------------------------------------------------------------------------

Created by: Sheryl Woods	Date: 3-30-08

Verified by:			 	Date:

Print Name

Sign Name