# SOP-E044

Beckman-Coulter Avanti- J-E Plate Centrifuge

Instructions for Use

Located in Bindley room 222

Purpose:  for the centrifugation of 384 well plates

Procedure:

1. tighten black round center control

2. cover plates with tape or covers

3. balance plates

4. set up the following settings:

a. ROTOR=   5.3

b. RPM(speed) = 400XG, press enter

c. TIME= 6min, press enter

d. TEMP= 4° C, press enter

5. Check settings:

a. show set, if correct, then press enter

b. press START, the lid will lock itself

6. When centrifugation is complete, the centrifuge will beep

7. Press the DOOR button to unlock lid.

-------------------------------------------------------------------------------------------------------------------------------

Created by: Kathy Ragheb,   April 27, 2011

Atoned by:

Verified by:			 	Date:

Print Name

Sign Name