SOP-E048

Use of VWR 91100 pH meter

INSTRUCTIONS

- Turn on meter (bottom left 2nd button)

- For calibration: you must remove old calibration

  - Hit CAL then CLEAR (CLR blinks in window)

  - Hit CLEAR a 2nd time – you see three bars

- Now insert electrode into pH 4 standard and wait for read

  - When the auto stops flashing, calibration for this is complete

- When 4.00 appears, clean electrode with wash and immerse in pH7 and hit READ

  - When the auto stops flashing, calibration for this is complete

- When 7.00  appears, clean electrode and wash and immerse in pH 10.01 and hit READ

  - When the auto stops flashing, calibration for this is complete

- Now the calibration is complete hit EXIT and the calibration is stored

- READ the Sample

- Clean electrode (wash) and Re-immers in electrode storage solution

NOTE: Never wipe electrode bulb

only use the water wash bottle to clean electrode