# SOP-S005

#### Summit 4.3 Uninstall/Re-Install Software

Objective:  To delete the Summit program files and re-install the operating software for the Beckman-Coulter CyAn ADP flow cytometer. All databases (.sum), protocols (.plo), and data files (.fcs) will not be deleted. Summit is verified to run on windows XP, service pack 2 only.

NOTE: a copy of the installation software is located on the D drive of the CyAn. (d:/service/swservice/xxxxxxxxx)

Procedure:

- If the system has not been totally shut down in a while, it would be a good idea to do so.

- shutdown the computer

- turn off the power to the CyAn (power switch is located on the back next to the power cable)

- turn off the power to the fluidics cart (power switch is located on the right side, near the power cable)

- wait a minute or two then turn everything back on in this order:

- Fluidics cart

- CyAn

- Computer

- You must uninstall the software and delete the registry files before re-installing the software

- If you are in Summit, exit out.

- go to START > control panel > add/remove programs

- scroll down to summit 4.3, click on it to select it

- click on the change/remove button

- select automatic > next > finish

- go to RUN > in the listbox type “regedit” > enter

- open the folder HKEY_CURRENT_USER > software > Beckman-Coulter

- delete the Beckman-Coulter folder

- open the folder HKEY_LOCOL_MACHINE . software > Beckman-Coulter

- delete this Beckman-Coulter folder

- close out of the registry editor

- install Summit 4.3 from either the disc, a flash drive, or drive D on the computer

- go to D:/service/swservice/xxxxxxxxxx

- if it is a zipped file, just right click on it and select option extract

- when it opens you will be welcomed to the Summit 4.3 Setup program

- select next

- select yes (License Agreement Screeen)

- Select Instrument type at down arrow  CyAn-ADP

- on site information screen type in West Lafayette, IN for the site, Purdue University for the Institutiponm, and Flow Lab for the Laboratory (or whatever)

- select next

- for Destination location select next

- for Set CyAn Model, slect LX with violet laser, next

- Set up type, typical, next

- Select Components Screen, next

- Backup Replaced Files, no, next

- Select Program Folder, next

- Start Installation, next

- When it is finished loading,  create a new database and Summit is ready to go.

---------------------------------------------------------------------------------------------------------------------------------

Created by: Kathy Ragheb 							Date: 05-13-03

Verified by:			 	Date:

Print Name

Sign Name