## Slide 1: Automated Reticulocyte Analysis: New Parameters for Anemia Diagnosis and Therapeutic Monitoring & Improved Precision & Laboratory Efficiency
Bruce H. Davis, M.D.
William Beaumont Hospital
Royal Oak, Michigan

## Slide 2: Automated Hematology: Desirable New Parameters
CD4 and CD8 Lymphocyte Subsets
Reticulocytes with immature reticulocyte maturation (IRF) - alias RMI
Neutrophil activation marker, such as quantitative PMN CD64 expression
reticulated platelets
Platelet activation markers (eg. CD62 or CD41 expression)
Hb F containing RBC enumeration (Kleihauer-Bettke)
Immune activation profile (cytokine/chemokine Rs)
CD5+ B Cells or light chain+ B cells (CLL, etc.))
Stem Cell enumeration (CD34+ cells)

## Slide 3: History of Automated Reticulocyte Counting
1980-90 Flow cytometric methods
  Tanke: Pyronin Y - Jacobberger: DiOC(3)
  Ortho: Acridine Orange - Others: PI, ethidium Br
  Metzger, Corash: Thioflavin T
  Lee (BDIS), Davis & Bigelow: Thiazole Orange
1990: TOA Sysmex R instruments: Auramine O
1992-96: Hematology instruments light scatter
  Technicon - H3: Oxazine 750
  Coulter Gen-S, STKS & MAXM: NMB
  Abbott Cell Dyn 3500: NMB
1996: Hematology instruments - fluorescence
  Abbott Cell Dyn 4000: thiazole-like dye
  Coulter Gen-S: CPO dye

## Slide 4: Automated Reticulocyte Counting: Methods Available - 1997
Thiazole Orange (BD) by Flow Cytometry
CPO dye (Coulter) by flow cytometry
TOA Sysmex R series and SE-Avante by Auramine O
Abbott Cell-Dyn 4000 by CD4K530
ABX Vega by thiazole orange
Bayer Technicon H3, Advia by oxazine dye
Coulter STKS/MAXM and Gen-S with new methylene blue (NMB)
Abbott Cell-Dyn 3500 with NMB
Fluoresence Methods
Light Scatter Methods

## Slide 5: Advantages of Automated Reticulocyte Analysis
Amenable to labor efficiencies or robotics
  faster analysis per sample
  allows for batch analysis or random access
Improved precision of retic counting
  superior to visual microscopic counts
  greater objectivity
New parameters of erythropoiesis
  Immature Reticulocyte Fraction (IRF)
  Reticulocyte hemoglobin content

## Slide 6: New Parameters with Automated Reticulocyte Analysis
Immature Reticulocyte Fraction (IRF)
  Reflects rate of erythropoietic activity
  Available on many instruments, methods
  Formerly termed reticulocyte maturity index (RMI)
  Replaces need for “corrected” reticulocyte count
Reticulocyte MCHC (hypochromic Retics)
  Detects early functional iron deficiency in Epo - Studies by Brugnaro, d’Onofrio
  Available only on Technicon H3 to date

## Slide 7
  IRF measured as fraction (0.00 - 1.00 range)
    Sysmex R: IRF = HFR + MFR (Ref Range: 0.05-0.20)
    Thiazole Orange: Cursor at 95% interval (Ref Range: 0.2-0.5)
  Report with reticulocyte % and absolute count
  Graphic display of retic count vs. IRF
    Superimpose refernce ranges for anemia classification
    Plotting sequential samples shows erythroid response
  Report results with other CBC parameters
  Flags for increased reticulocytosis and hypoproliferative response
  Automated, random access, discrete testing
Reticulocyte Enumeration with
Immature Reticulocyte Fraction (IRF)

## Slide 8: Immature Reticulocyte Fraction (IRF) Thiazole Orange by Flow Cytometry
IRF = #HFR/#Retics
Data Analysis
  exclude nucleated cells (nRBCs, PMNS, lymphs)
  exclude platelets
  define IRF region
  define retics
RBCs
Reticulocytes
Platelets
Nucleated
Cells
IRF

## Slide 9: Evidence for Pathophysiologic Relevance of Immature Reticulocyte Fraction (IRF)
Erythropoietin therapeutic effect: IRF 1- 3 days
  CD71 vs TO studies
  BJH study, Major et al.
Animal models
  in vivo biotinylation studies
  CD71 vs. TO studies
BMT recovery
  IRF earliest parameter of engraftment
  various methods with demonstrated efficacy

## Slide 10: Normal Erythopoiesis
Maturational continuum
EPO effect
blood retic populations
  IRF retics (CD71+)
  late retics
  stress retics

## Slide 11: Reticulocyte Maturation: in vivo biotinylation (K. Ault)
72 hours
6 hours
24 hours
pre-biotinylation

## Slide 12: Bone Marrow Regeneration Response: Consistent Pattern

## Slide 13: Erythroid Parameters with Erythropoietic Response

## Slide 14: Evaluation of Erythropoiesis: Bivariate IRF and Retic Count Display

## Slide 15: Immature Reticulocyte Fraction (IRF): Clinical Utility in Medical Practice
Monitor BM or Stem Cell Regeneration post-BMT or ChemoRx
Monitor Renal Transplant Engraftment (Epo production)
Monitor Neonatal Transfusion Needs
Monitor Anemia Therapy
Monitor EPO Therapy: Renal Failure, AIDS, Infants, MDS
Monitor Bone Marrow Toxic Insults from drugs (eg. AZT)
Prognostic in Anemia of AIDS and Prematurity
Timing for Stem Cell Harvests following Growth Factor or Cytotoxic Drug Therapy
Detection of Aplastic Crisis in Hemolytic Anemias
Diagnosis and monitoring of aplastic anemia
Evaluate Normochromic Anemias of Various Etiologies
Detection of Occult or Compensated Hemorrhage or Hemolysis
Classification of Anemias

## Slide 16: Patterns of IRF and Retic counts in Anemia
Aplastic anemia/crisis
hypoplastic anemia
BM regeneration
Chronic disease
Iron deficiency
Thalassemia
Folate/B12 deficiency
Myelodysplasia
Hemolytic anemia
Blood loss/anoxia
Low
Low
Low
Low/WNL
Low/WNL
WNL/high
Low/WNL
Any level
High
WNL/high
Low
Low
High/WNL
WNL
High
WNL/high
High
WNL/high
High
High
Clinical Condition
Retic Ct
IRF

## Slide 17: Intermethod Correlation Studies
Single site studies: multiple published
  improved precision - CVs <15%
  intermethod bias, but “clinically insignificant”
Davis et al: AJCP 102:468, 1994
  8 sites, 11 instruments, 310 blood samples
  IRF and Retic counts compared
College of American Pathologists’ Reticulocyte RT Survey 1994-96
  >2,600 participants
  surrogate blood material
  Retic % only reported

## Slide 18: Retic Counts: Inter-method Correlation

## Slide 19: IRF: Inter-method Correlation

## Slide 20: Reticulocyte Proficiency Testing College of American Pathologists
1994-96 Program - surrogate blood material
Methodologies Approved for Testing
New Methylene Blue visual microscopy
Sysmex R series (auramine O)
Flow Cytometry - Thiazole Orange
Flow Cytometry - Other dyes
Coulter STKS/MAXM - NMB
Miles Technicon H3 - Oxazine

## Slide 21
CAP Survey 1995 RT-C: Distribution of Methods

## Slide 22
<footer>
<date/time>

## Slide 23
<footer>
<date/time>

## Slide 24
A B C D E F G
Flow Cytometer Instrument
1 2 3 4 5 6
Specimen Number
A
B
CAP RT Survey Experience: No Bias with TO methods
secondary to FCM instrument

## Slide 25: Known or Potential Interferents
Cellular Elements
  Platelet clumps or giant platelets
  nucleated cells or fragments
RBC Inclusions
  Howell-Jolly bodies
  Heinz or Pappenheimer bodies
  parasites (malaris, babesia)
Miscellaneous causes
  Autoflourescence (drugs, porphyria)
  RBC aggregation (paraproteins, cold agglutinins)
  coincidence (eg. platelet and RBC)
  abnormal RBCs, hemolysis

## Slide 26: Controls for Clinical Practice
Commericial Preparations
  R&D Systems, Minneapolis, Mn
  Streck Lab, Omaha, Ne
  Instrument manufacturers
Refrigerated blood samples
  short term QC by carry-over comparison
  least expensive
  will not detect long-term drift
Veterinary blood samples
  rabbit
  porcine

## Slide 27: Reasons for NOT utilizing automated reticulocyte counting
Volume does not exceed 3-5/day
Physicians expect “stat” results
Waiting for the “next generation” instrument
“We’ve always done it this way”
Technologists like doing manual counts