## Slide 1: Field-Deployable Laser-Induced-Breakdown-Spectroscopy for Multiplexed Detection of Food Contaminants Using Rare-Earth Labeled Antibodies Embedded in a Lateral Flow Strip
J. Paul Robinson1,2 Carmen Gondhalekar 1,2, Bartek Rajwa1,3, Euiwon Bae4, Iyll Joon Doh, Valery Patsekin1, Larry Stanker5
1Basic Medical Sciences, College of Veterinary Medicine; 2Weldon School of Biomedical Engineering; 3Bindley Bioscience Center, Discovery Park, 4School of Mechanical Engineering; Purdue University; 5USDA-ARS, CA

## Slide 2: Platform Requirements for Toxin and Pathogen Detection
Identifying pathogenic organisms requires either a highly specific probe or label-free detector capturing the structure/chemistry or phenotype.
If there is no/little background the accuracy, specificity, and resolution can be high
Multiplexing leads to cost reduction/efficiency
Rapid analysis is mandatory for this technology to be acceptable

## Slide 3: Toxins of Interest to this project
Shiga toxin are from the bacteria Shigella dysenteriae and the shigatoxigenic serotypes of Escherichia coli (STEC), which includes serotypes O157:H7, O104:H4, and other enterohemorrhagic E. coli (EHEC).
Botulinum neurotoxins (BoNTs) are the causative agents of the deadly food poisoning disease botulism, and could pose a major biological warfare threat due to their extreme toxicity and ease of production. They also serve as powerful tools to treat an ever expanding list of medical conditions.
Ricin is a highly toxic, naturally occurring lectin (a carbohydrate-binding protein) produced in the seeds of the castor oil plant, Ricinus communis. A dose of purified ricin powder the size of a few grains of table salt can kill an adult human.The median lethal dose (LD50) of ricin is around 22 micrograms per kilogram of body weight if the exposure is from injection or inhalation (1.78 milligrams for an average adult). Oral exposure to ricin is far less toxic as some of the poison is inactivated in the stomach. An estimated lethal oral dose in humans is approximately 1 milligram per kilogram.

## Slide 4: Multiplexed Detection of Biological and Chemical Contaminants Current bio-contaminant detection strategies are time and resource intensive – but also very sensitive
  Toxin Lethality
  Botulinum: 2.9 µg /kg body weight
  Shiga: 2.9 µg /kg body weight
  Ricin: 1-20 mg /kg body weight
  Assays for toxin detection
  Mouse lethality – 4 days
  Stool culture – 2-3 days
  Enzyme immunoassay, Vero cell cytotoxicity assay & PCR – 1 day
  Assay Sensitivity
  Botulinum detection: mouse lethality: 20-30 pg/ml, Immuno-PCR: 5pg/ml
  Shiga toxin detection: commercial EIA 10 ng/ml, Immuno-PCR: 1pg/ml
  Ricin toxin detection: immuno-PCR: 1pg/ml
Once developed, the technology can be multiplexed for many different antigens of interest in a multiplex of around 10-15 agents
Notes: Conventional toxin identification and quantification methods are very sensitive but also time consuming.
On this slide I list some common FDA-approved toxin identification strategies, and you will notice that the shortest assay takes at least a day to complete.
These assays have been good for toxin detection because the amount of toxins they can detect is well below the amount that it takes for the toxin to be lethal
Here you can see and compare some values for botulinum, shiga and ricin

## Slide 5: Spark-induced breakdown spectroscopy (SIBS) & Laser induced Breakdown Spectroscopy (LIBS) utilize ionic plasma to characterize materials
  Techniques that utilize the 4th state of mater to determine sample composition
  CyTOF, ICP, LIBS, SIBS
HEAT
Chart: 
Notes: The spectra displayed is a SIBS spectra
SIBS and LIBS are very similar – the technique excites atoms to a higher energy state by displacing electrons to higher energy orbitals (or from the atom completely) using a pulse of energy – either a strong electrical potential or electromagnetic beam. As the atoms return to resting state, the transitioning electrons release energy as light. The wavelength of emitted light corresponds to the transition energy. Since each element has a unique assembly of electrons and orbitals, the spectra of emitted wavelengths and their intensities is unique for each element. Therefore the elemental composition of a material can be determined by its spectral emission.
Possible questions:
How is LIBS and SIBS emissions different than fluorescence?
Why are you advancing from SIBS to LIBS?

## Slide 6: MARS “Curiosity” Rover since 2012
LIBS
Images From
http://appliedspectra.com/mars-libs.html

## Slide 7
Images From
http://appliedspectra.com/mars-libs.html
Laser damage from the pulsed lasers on mars rock!

## Slide 8: Rationale for selecting Label-based detection
Fluorophores
Rare Earth Metals
…SIBS or LIBS
CyTOF
Application in Hematology
Immunology
Major use is phenotyping
blood cells - bone marrow
Cell lines, etc.
It is based on using
lanthanide-antibody tags at
very high multiplexing capacity
Notes: Emphasize the type of
Technology you are focusing on
Emphasize differences
Don’t forget field deployable

## Slide 9
Figure from: http://www.the-scientist.com/?articles.view/articleNo/31438/title/Flow-Cytometry-for-the-Masses/

## Slide 10
First rare earth element found about 1800 being yttrium and cerium
40,000 tonnes mined each year – 80% from China
So not really rare but were though to be when isolated and kept the name…
Seracon poster - Keith Enevoldsen elements.wlonk.com
Rare Earth Elements
Notes: Possible question: So if they are not rare, then why do people use these metals?
Possible question: Because they are common in so many things, are you worried that food samples (which is processed on machinery where these materials may exist), will create false positives?

## Slide 11
http://www.hematology.org/Thehematologist/Mini-Review/4553.aspx
https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4945403/pdf/nihms-760310.pdf
Currently, gadolinium-containing complexes – gadopentetic acid (Magnevist®) and gadoteric acid (Artirem®) – are commonly used as MRI contrast agents for cancer imaging, while lanthanide radioisotopes like 177Lu have been used in cancer imaging and therapy.

## Slide 12: Our approach

## Slide 13: Lanthanide Conjugation Chemistry
Antibody conjugation to heavy metal polymers
  An acrylic polymer is formed with a reproducible number of active NHS-groups. Chelators are then bound to the NHS-groups.
  Metal isotopes bind to the chelators
  A reactive thiol group is added to the end of the polymer – this group binds to the disulfide bond in the Fc region of the Ab
  There are 2-3 polymers per antibody; 30 metal atoms per polymer
Antibodies conjugation to metal nanoparticles
  NHS-active or COOH coated nanoparticles react with primary amines on antibodies, binding them to the surface of the particle
Capture antibody:
Anti-E. coli
Anti-botulinum
Anti-shiga toxin
Anti-ricin

## Slide 14: Multiplexing – the big picture concept Potential long term opportunities
Rare metal | Cat A Panel 1 | Cat A Panel 2 | Cat B Panel 1 | Cat B Panel 2 | Cat B Panel 3 | Cat B Panel 4 | Cat C Panel 1
Er | Bacillus anthracis |  |  |  | Cryptosporidium parvum | Prions | Nipah virus and additional hantaviruses
Nd | Clostridium botulinum |  |  |  | Cyclospora cayatanensis | Other Rickettsias | Tickborne hemorrhagic fever viruses
Ho | Yersinia pestis |  | Coxiella |  | Giardia lamblia | Rabies | Crimean-Congo Hemorrhagic fever virus
Tb | Variola major |  | Brucella species | Diarrheagenic E. coli | Entamoeba histolytica | West Nile Virus | Tickborne encephalitis viruses
Gd | Francisella tularensis |  | Burkholderia mallei | Pathogenic Vibrios | Toxoplasma | LaCrosse | Yellow fever
Tm |  | Hantaviruses | Chlamydia psittaci | Shigella species |  | California encephalitis | Tuberculosis, incl drug-resistant TB
Pr |  | Rift Valley Fever | Epsilon toxin C.perfringens | Salmonella |  | Venezuelan equine encephalitis | Influenza
Sm |  | Dengue | Staph ent’tox B | Listeria monocytogenes |  | Eastern equine encephalitis | Chikungunya virus
Eu |  | Ebola | Rickettsia prowazekii | Campylobacter jejuni |  | Western equine encephalitis | SARS-CoV
Yb |  | Marburg |  | Yersinia enterocolitica |  | Japanese Encephalitis Virus | Coccidioides immitis
La |  |  |  | Diarrheagenic E.coli |  | Kyasanur Forest Virus | Coccidioides posadasii

## Slide 15: Multiple level Panels
Listeria monocytogenes
Label-free pattern
Diarrheagenic E.coli
Bacilus anthracis
Category A
Panel 1
Category A
Panel 2
Category A
Panel 3
High-level panel
Low-level panel
Phase 1
Phase 2
Phase 3
Initial Hardware design

## Slide 16: SIBS Instrumentation is designed to induce a plasma and collect the electromagnetic emission
Notes: Emphasize that we don’t want to collect the spark, only the plasma

## Slide 17: Creating calibration standards and sample delivery systems
Chart: Au peak intensity vs coating time
| 0 | 1 | 2 | 3
Spark 1 | 62.375 | 655.619 | 2940.5833 | 5298.9167
Chart: 
Calibration standard:
Silica wafer
Gold coated silica wafer
Copper foil
Sample holder:
Silicone wafer
Whatman paper
Nylon mesh
Porous silicon
Porous polystyrene
1 min
Silicon wafer
Gold
nano-particles
Gold-coated silicon wafer
0 min
2 min
3 min

## Slide 18: Electronic triggers must be perfectly timed to collect emissions
Chart: Atomic Spectra: Copper foil Delay time = 12.14 µs
Chart: Atomic Spectra: Copper foil Delay time = 2.9 µs
Notes: Talk about timing between trigger and collection
Can relate timing to the time it takes for cell detection in flow cytometry/cell sorting

## Slide 19
Labeling confirmatory study – Energy dispersive X-ray spectroscopy (EDX)
Gd on wafer
Pr on wafer

## Slide 20: Confirmation of Data with LIBS
Laser
Spectrometer
Sample

## Slide 21: Small benchtop LIBS to test optics and lasers
Helium tank
Q-switched
laser
Sample

## Slide 22: Preliminary studies indicate that heavy-metals tags on anti-toxin antibodies can be detected using LIBS & SIBS
Dy II
Pr II
Si I
Nd:YAG
Laser
Pr
Silicone wafer
Laser induced breakdown spectroscopy uses a laser to
Induce a plasma instead of a spark
0.6mg Pr-Antibody
Notes: we started off with libs- it was the concept basis of the project
LIBS gave us excellent preliminary data but it’s still expensive and the cost is going down
SIBS is more sensitive, less expensive, and more field deployable

## Slide 23
NIST database: Au Atomic Spectra
Experimental: Au Atomic Spectra
NIST database: Gd Atomic Spectra
NIST – National Institute of Standards and Technology
Overlapping NIST database with experimental spectra
permits peak identification
Peaks are based on data from the NIST database
Notes: Mention hannover and making our own database
Make sure you know what observed vs. ritz is

## Slide 24
Two methods for labeling antibodies with metals
Paper-based diagnostics
Dy
Eu
Yb
Yb
Eu
Dy

## Slide 25
Paper-based diagnostics
Integrating metal-labeled antibodies in paper bioassays
Using different metals in paper-based assays would create great multiplexing potential
Metal
X
Metal X

## Slide 26
Paper-based diagnostics
Integrating metal-labeled antibodies in paper bioassays
Each metal label can be used to indicate the presence of a certain antigen
Dy
Anti-botulinum
Eu
Anti-shiga toxin
Lu
Anti-ricin
Yb
Anti-aflatoxin
Dy
Eu
Lu
Yb
Yb
Eu
Dy
Lu
Notes: Improve automation
Maybe include little stars to indicate animations.

## Slide 27
Yb
Eu
Dy
Lu
Paper-based diagnostics
Integrating metal-labeled antibodies in paper bioassays
Using metal nanoparticles other than gold and silver for paper assays is uncommon – why?
Dy
Anti-botulinum
Eu
Anti-shiga toxin
Lu
Anti-ricin
Yb
Anti-aflatoxin
?
Dy
Eu
Lu
Yb
Anti-aflatoxin
Anti-ricin
Anti-botulinum
Notes: Improve automation
Maybe include little stars to indicate animations.

## Slide 28
Laser-induced breakdown spectroscopy (LIBS)
Shiga toxin (Eu label):
Botulinum (Lu label):
Ricin (Dy label):
Aflatoxin (Pr label):
Using different nanoparticles in a single bioassay opens the door to multiplexing
Dy
Anti-botulinum
Eu
Anti-shiga
Lu
Anti-ricin
Yb
Anti-aflatoxin
Laser-induced breakdown spectroscopy: a tool for elemental analysis of materials
US Patent US20160161415 A1 (US 15/042,969)
Notes: No one has translated portable LIBS to the biological realm
Purdue has a patent application. Include patent application number.

## Slide 29: A simple LIBS system detected metal-labeled antibodies
Laser-induced breakdown spectroscopy (LIBS)
Eu-labeled antibodies spectra
Europium peaks detected
Ca
Fe
Na
Sr
Cu
Air
Notes: Europium emits a bluish light

## Slide 30
Anti-E. coli
Yb
Au
Laser-induced breakdown spectroscopy (LIBS)
Eu
Dy
Pilot study detects metal-labeled antibodies on paper
US Patent US20160161415 A1 (US 15/042,969)

## Slide 31
Laser-induced breakdown spectroscopy (LIBS)
A simple LIBS system detected metal-labeled antibodies
Yb
Eu
Dy
Au

## Slide 32
LIBS results were validated using LA-ICP-OES
Chart: Detection of Dy Abs and ICP standards on nitrocellulose
| Dy (340.780) | Dy (353.171)
Dy Abs | 5010.6045 | 13124.9876
Dy ICP | 110420.3378 | 285773.961
Control | 22.3272 | 42.1698
Chart: Detection of Eu Abs and ICP standards on nitrocellulose
| Eu (397.197) | Eu (420.504)
Eu Abs | 48565.8529 | 71768.3633
Eu ICP | 2694361.9062 | 4323536.5625
Control | 262.446 | 469.6025
Chart: Detection of Au nanoparticles on nitrocellulose
| Au (208.207) | Au (242.794) | Au (267.594)
Au line | 1278.3474 | 5674.9759 | 8852.9011
Control | 23.4315 | 59.4503 | 91.3855
(Inductively coupled plasma atomic emission spectroscopy)

## Slide 33
Room Size
Bench Size
Portable
Research Lab
Testing Lab
Field

## Slide 34: Portable hardware must be compact There have been major tech developments recently
Small
Light
Efficient
Sensitive
Low Power draw
Battery Operated
Fast
(Cost)
Need at least 2 spectrometers
Spectrometer
Q Switched Laser
Precious metals
Geology

## Slide 35
Testing tiny powerful lasers
500 Hz, varied pulse number
Top left: 10 pulses, 500hz
Top right: 1000 pulses, 500hz
1000 Hz, varied pulse number
Bottom left: 200 pulses, 1000hz
Bottom middle: 500 pulses, 1000hz
Bottom right: 1000 pulses, 1000hz

## Slide 36: Handheld – Geotagged Instrument
Chemistry
Sampling method
Analysis
Detectiony

## Slide 37
Multiplexed detection of biological and chemical food contaminants using spark-induced breakdown spectroscopy (SIBS)
Solid-phase immuno-assay
1) Toxin extraction
2) Immobilization
3) Labeling
Instrumentation
Electronics
Optics
Timing
Data Analysis
Software design
Spectral library
Spectral matching
Conclusions
Ln
Gd
Eu
Pr
Detect lanthanide-tagged antibodies with SIBS and CyTOF
Multiplexing for immuno-tags
Measuring chemicals
Bench based
Portable
Rapid
Definitive
Notes: Field deployable

## Slide 38: Acknowledgements
  Dr. Eva Biela
  Carmen Gondhalekar
  Kathy Ragheb
  Jennifer Sturgis
  Dr. Valery Patsekin
  Dr. Euiwon Bae
  Iyll-Joon Doh
  Dr. Huisung Kim
  Dr. Bartek Rajwa
  Dr. Larry Stanker (UDSA)
  Dr. Prasoon Diwakar
  Willy Costakis
This material is based upon work supported by the NIH Grant# 1R56AI08951-01 and the U.S. Department of Agriculture, Agricultural Research Service, under Agreement #.1935-42000-072-02G to the Center for Food Safety Engineering. Any opinions, findings, conclusion, or recommendations expressed in this publication are those of the author(s) and do not necessarily reflect the view of the U.S. Department of Agriculture. CG is a NSF Fellow (Grant# DGE-1333468.