## Slide 1: Advances in Cell Analysis J. Paul Robinson, SVM Professor of Cytomics, Professor of Biomedical Engineering, Purdue University, USA
Sysmex Presentation
June 2014

## Slide 2: Purdue University Cytometry Laboratories Acknowledgements
Contributors
  Staff Scientists: Valery Patsekin,
Graduate Students: Awais Mansoor, Angela Wolf, Fiona Thomas, Anthony Pedley
Undergrad Students: Nicole Lewis, Peter Carlsgaard
Faculty: Bartek Rajwa, Euiwon Bae, J. Paul Robinson
Data sets: Gary Nolan Lab, David Hedley, Bernd Bodenmiller, Paul Wallace
Staff
  Jennie Sturgis (Imaging)
  Kathy Ragheb (Flow)
  Cheryl Holdman (Flow)
  Raymond Fatig (cell culture)
Funding:
NIH, NSF, USDA, Purdue University
Corporate: Beckman-Coulter, Point-Source, Parker-Hannifin, Polysciences, Bangs labs, Roche, MediaCybernetics, Q-Imaging, Kodak Medical Systems, Crystalplex, Becton-Dickinson , Icyt, eBioscience,
Edmund Optic, Digilab Inc., Spherotech, Inc..
IntelliCyt, Inc
Research reported in this presentation was supported by the National Institute of General Medical Sciences & the National Cancer Institute of the National Institutes of Health under Award Numbers 1R56AI089511-01, R21RR033593-01 & 1R33CA140084. and the UDSA under award # ARS 342365. The content is solely the responsibility of the authors and does not necessarily represent the official views of the National Institutes of Health.
Bindley Bioscience Center, Purdue University

## Slide 3: Background and Concepts
Flow cytometry saw little changed over its first 40 years:
  An historical Review
Some emerging technologies:
  Spectral cytometry
  HT flow cytometry
  The impact of HT cytometry
  The impact of next generation analytical concepts
Part 1
Part 2
Part 3

## Slide 4: Tyndall 1869
1869

## Slide 5: Andrew Moldavan
It is unclear if Moldavan ever built his cell counter. His short article (less than half a page) describes a number of problems but no results.
Manuscript:
Photo-Electric Technique for the Counting of Microscopical Cells
Andrew Moldavan
Montreal, Canada
Science 80:188-189, 1934
“The purpose of the experiment is to have each microscopical cell passing through the capillary tube, register itself automatically on the photoelectric apparatus, thus creating a micro-current which can be amplified and recorded.”
Moldavan (1934) demonstrates use of a suspending fluid in which were blood cells - the measurements were made in a capillary tube using a photoelectric sensor to make extinction measurements
1934

## Slide 6
1934

## Slide 7: Conjugated Fluorescein with Isocyanate
Coons and Kaplan (1950) - conjugated fluorescein with isocyanate - better blue green fluorescent signal - further away from tissue autofluorescence. This method used a very dangerous preparative step using phosgene gas
1950

## Slide 8: Gucker’s Apparatus
A Photoelectronic Counter for Colloidal Particles1
BY FRANK T. GUCKER, J R . , ~ CHESTER T. O’KONSKI,~ HUGH B. PICKARD3 AND JAMES N. PITTS, JR.4
1947
Developed a flow cytometer for detection of bacteria in aerosols
Instrument: Sheath of filtered air flowing through a dark-field flow illuminated chamber. Light source was a Ford headlamp, PMT detector (very early use of PMT)

## Slide 9: Gucker’s Apparatus
A Photoelectronic Counter for Colloidal Particles1
BY FRANK T. GUCKER, J R . , ~ CHESTER T. O’KONSKI,~ HUGH B. PICKARD3 AND JAMES N. PITTS, JR.4
1947
Developed a flow cytometer for detection of bacteria in aerosols
Instrument: Sheath of filtered air flowing through a dark-field flow illuminated chamber. Light source was a Ford headlamp, PMT detector (very early use of PMT)

## Slide 10: Wallace Coulter
Wallace Coulter - Coulter orifice - 1956 -
(patent 1953) - measured changes in electrical conductance as cells suspended in saline passed through a small orifice
The first commercial version of the Coulter Counter
Coulter’s
Original 1953
Patent app’n
From 1949
©J.Paul Robinson
1949
1949 - Coulter WH, US Patent 2,6756, 508.
“Electrical sensing zone method”)

## Slide 11: Conjugated Fluorescein with Isocyanate
Coons and Kaplan (1950) - conjugated fluorescein with isocyanate - better blue green fluorescent signal - further away from tissue autofluorescence. This method used a very dangerous preparative step using phosgene gas
1950

## Slide 12: Early Microfluorometric Scanner Robert Mellors 1951
RC Mellors & R. Silver, A microfluorometric scanner for the differential detection of cells: application to exfoliative cytology, Science 104, 1951
1951

## Slide 13: Marylou Ingram
1951

## Slide 14: P.J. Crosland-Taylor
Sheath Flow Principle – 1952-3
A Device for Counting Small Particles Suspended in a Fluid through a Tube
P.J. Crosland-Taylor
Bland-Sutton Institute of Pathology
Middlesex Hospital, London, W.1. June 17, 1952 Nature 171: 37-38, 1953
“Provided there is no turbulence, the wide column of particles will then be accelerated to form a narrow column surrounded by fluid of the same refractive index which in turn is enclosed in a tube which will not interfere with observation of its axial content.”
1952

## Slide 15
1952
Marylou Ingram in medical school 1944
1960

## Slide 16: Fundamental principle of Sheath Flow PJ Crosland-Taylor
A device for counting small particles suspended in fluid through a tube, Nature 171, 1953 (ref 4756)
An Electronic Blood-Cell Counting Machine, Blood 13, 1958 (ref 40974)
1953

## Slide 17: 1500 Coulter Counters
1961

## Slide 18
1953

## Slide 19: Before Cytometry
Slide Kindly Supplied by Compucyte
Lou Kamentsky – at the IBM’s Watson Labs 1950s and 1960s
1960s

## Slide 20
1962
Valves, Valves and Valves….

## Slide 21
1962

## Slide 22
1962

## Slide 23
1963

## Slide 24: TOA Enters blood analysis field - 1963
1963-Automatic blood cell counting instrument developed and first unit delivered. CC-1001 developed
Sysmex Journal International, Vol 9, #1, 1999

## Slide 25: Historical Overview
Early flow systems
Hallermann et al, Kosenow - 1964 - AO staining of leukocytes - was able to use fluorescence (in a flow based system) to select leukocytes from red cells despite a low ratio (1/1000) because they took up lots of AO - also claimed to be able to discriminate between monocytes and PMN
Scan and structure from www.probes.com
Molecular Formula: C17H20ClN3
Molecular Weight: 301.82
CAS Number/Name: 65-61-2 / 3,6-Acridinediamine, N,N,N',N'-tetramethyl-, monohydrochloride
1964
Halleman 1964 AO

## Slide 26: Richard Sweet
Richard Sweet developed the electrostatic ink-jet printer which was the principle used by Mack Fulwyler to create a cell-sorter.
1964

## Slide 27: Electrostatic Printing and Droplet Sorting RG Sweet - MJ Fulwyler
RG Sweet, Fluid Droplet Recorder
U.S. Patent 3,596,275
filed 3/25/64 (ref 40975)
MJ Fulwyler, Particle Separator, U.S.
Patent 3,380,584 filed 6/4/65
1965

## Slide 28: Mack Fulwyler
Mack Fulwyler – worked in Marvin van Dilla’s lab at Los Alamos. – developed the sorter in 1965 – initially used electronic cell volume - at Los Alamos National Labs - this instrument separated cells based on electronic cell volume (same principle as the Coulter counter) and used electrostatic deflection to sort. The cells sorted were RBC because they observed a bimodal distribution of cell volume when counting cells - the sorting principle was based on that developed for the inkjet printer by Richard Sweet at Stanford in 1965.
After determining that the bimodal distribution was artifactual, this group were able to sort neutrophils and lymphocytes from blood.
Marvin van Dilla
(Photo dated 6/91)
Mack Fulwyler
Lushbaugh, CC, Basmann,NJ, Glascock, B. Electronic measurement of cellular volumes.
II Frequency distribution of erythrocyte volume. Blood 1962 20:241-248 Correspondence Blood 1964:23:403-405
1965

## Slide 29: Sorting Cancer Cells - Kamentsky
Spectrophotometric Cell Sorter, LA Kamentsky and MR Melamed, Science 156, 1967 (ref 4134)
1965
Technology was based on a fluidic switched “Typewriter” that IBM was developing in Zurich research laboratories.

## Slide 30: Four Sensors, Sorting, Auto Sampling and Computer Data Reduction
Two analytic instruments were built and one was delivered to LA Herzenberg at Stanford University 1967
IBM 1620 Computer
1966
This instrument became the “Hemalog D” by Technicon in 1974

## Slide 31: TOA Automatic Hematology Analyzer - 1966
Automated Hematology Analyzer, CC-1002

## Slide 32
1967

## Slide 33
1968

## Slide 34
J.Pharm.Sci 57:557-563, 1968
1968

## Slide 35: Sysmex precursor formed - 1968
1968 - -Establishes a medical equipment sales company, Toa Medical Electronics Co., Ltd. (now Sysmex Corporation).

## Slide 36: Dittrich & Göhde
Dittrich & Gohde - 1969 - Impulscytophotometer (ICP)- used ethidium bromide for a DNA stain and a high NA objective used as a condenser and collection lens
Phywe AG of Gottingen (1969) - produced a commercial version of the ICP built around a Zeiss fluorescent microscope
Laerum, Göhde, Darzynkiewicz (1998)
Photos ©2000 – J.P. Robinson
Göhde and Laerum (1998)
1969
http://www.partec.de/partec/flowmuseum.html
ICP 11 (1969)
Distributed by Phywe, Göttingen
The first commercial flow cytometer
PDP 11 computer

## Slide 37: Herzenberg Lab
Len Herzenberg - 1969 - sorter based on fluorescence (arc lamp) built after working with one of Kamentsky’s RCS systems where they built an instrument they called the Fluorescence Activated Cell Sorter (FACS)
Photos from 2000 – JP Robinson
1969

## Slide 38: Kamentsky & Biophysics Systems
Kamensky’s first benchtop instrument the Cytograph. This measured scatter using a He-Ne laser. This particular instrument was a model prior to the fluorescence detection model.
Kamentsky also built a fluidic cell-sorter to evaluate the cells identified in his RCS. An RCS was sent to Stanford for use by Leonard Herzenberg . The unit was also the model for the Technicon D instrument built by Technicon. (Reference: Shapiro)
©J.Paul Robinson
1970

## Slide 39: Evolution of Leukocyte Gating Strategy for Cluster Subsetting
GC Salzman, JM Crowell, JC Martin, TT Trujillo, A. Romero, PF Mullaney, & PM LaBauve,
Cell classification by laser light scattering: identification and separation of unstained leukocytes,
Acta Cytologica 19, 1975
LR Adams & LA Kamentsky,
Machine characterization of human leukocytes by acridine orange fluorescence, Acta Cytologica 15, 1971
RA Hoffman, PC Kung, WP Hansen,
Simple & rapid measurement of human T
lymphocytes and their subclasses,
PNAS 77, 1980
Lymphocytes
Monocytes
Granulocytes
1971
Biophysics Systems
Los Alamos
Ortho

## Slide 40: Herzenberg & Becton-Dickinson
Herzenberg -1972 -2013 Argon laser flow sorter - placed an argon laser onto their sorter and successfully did high speed sorting - Coined the term Fluorescence Activated Cell Sorting (FACS) This instrument could detect weak fluorescence with rhodamine and fluorescein tagged antibodies. A commercial version was distributed by B-D in 1974 and could collect forward scatter and fluorescence above 530 nm.
Herzenberg was the recipient of the very
Prestigious 2006 Kyoto Prize for his work
in development of fluorescence based
flow cytometry. Many people well known
in the field were trained in his lab
(Photo from the official Kyoto Prize website)
1972

## Slide 41: Herzenberg and Bonner
WA Bonner, HR Hulett, RG Sweet and LA Herzenberg,
Fluorescence Activated Cell Sorting, Review of Scientific Instruments 43, 1972
1972

## Slide 42: TOA 1972
1972 - -Transfers medical equipment division operations to
Toa Medical Electronics Co., Ltd. (now Sysmex Corporation).

## Slide 43
Biophysics Systems
Ortho Instruments
Cytograf (1970)
Cytofluorograf (1970)
FC200 (1975)
ELT 8 (1977)
50H (1978)
Spectrum (1979)

## Slide 44: BD
1974
Coulter
FASEB April 1974 – Coulter Booth

## Slide 45: EPICS II
1974
Delivered to Chet Herman
NIH in 1974

## Slide 46: Hemalog D
Technicon - Hemalog D - 1974 - first commercial differential flow cytometer - light scatter and absorption at different wavelengths - chromogenic enzyme substrates were used to identify neutrophils and eosinophils by peroxidase and monocytes by esterase, basophils were identified by the presence of glycosaminoglycans using Alcian Blue - the excitation for all measurements was a tungsten-halogen lamp
Insert photos on page 60
Image from Shapiro “Practical Flow Cytometry”, 3rd. Ed.Wiley-Liss, 1994
1974

## Slide 47: Mack Fulwyler
Coulter Electronics manufactured the TPS-1 (Two parameter sorter) in 1975 which could measure forward scatter and fluorescence using a 35mW argon laser.
This photo (left) is one of only one or two surviving TPS Instruments. It is very similar to the Coulter Counter of the day.
Photo ©2000 – J.P. Robinson
1975

## Slide 48: Monoclonal Antibodies
Georges J.F. Köhler
César Milstein
1975

## Slide 49
1975

## Slide 50: TOA Releases Automated Hematology Analyzer
1975 - Released the first fully Automated Hematology Analyzer developed in Japan, "CC-710.“

## Slide 51: Biophysics -Ortho
Ortho Diagnostics (Johnson and Johnson) purchased Biophysics in 1976 and in 1977 the System 50 Cytofluorograph was developed - this was a droplet sorter, with a flat sided flow cell, forward and orthogonal scatter, extinction, 2 fluorescence parameters, multibeam excitation, computer analysis option.
1979 - NIH scientists had added a krypton laser at 568 nm to excite Texas Red fluorescence at 568 nm and emit at 590-630 nm. Argon (488 nm FITC was measured simultaneously without signal cross-talk - thus the FACS IV was developed (B-D).
FC 200(1975)
System 50H (1978)
1977

## Slide 52: Chromosome Analysis on the TPS-1 Sorter
1977

## Slide 53: Two groups of Lymphocytes B Cells and T Cells….
1977

## Slide 54: Coulter Electronics
1977-78 developed the Epics series of instruments which were essentially 5 watt argon ion laser instruments, complete with a multiparameter data analysis system, floppy drive and graphics printer.
Epics V front end (left) and MDADS (right)
Photo ©2000 – J.P. Robinson
1978

## Slide 55: 1978 Sysmex corporation born

## Slide 56: And so we get to the present…
Millipore/Guava
Apogee
Luminex
Accuri
Bryte
BC
BD
BD
BC
BC
BC
Amnis
Bay
Biosciences
BD Influx
BD Aria
Sony
BC
BC

## Slide 57: The next generation of cytometry…
Spectral Cytometry
Part 2

## Slide 58: Hyperspectral Single Cell Cytometry
10 μs available
10,000-100,000 cells
64 ch ADC`
Built in 2002-2003

## Slide 59: Detectors
32 spectral channels Hamamatsu
6 regular fluorescence channels
1 side scatter channel
6 forward scatter detectors
1 time gate
ADC Electronics
64 ch 16 Bit ADC
Purdue Issued Patent #7,280,204

## Slide 60
PE
Rh-123
Chlorophyll
Bang’s
Labs
Dyed
Beads*
These CANNOT be resolved
Advanced classification tools
like PCA are required
“2” color analysis
These CAN now be resolved

## Slide 61: Spectral “fingerprints” identify samples

## Slide 62: Spectral “fingerprints” identify samples
This technology is now a commercial technology (Sony) (…to HT)

## Slide 64
5-(and-6)-carboxy-2´,7´-dichlorofluorescein diacetate (CDCFA)
5(6)-carboxy-4',5'-dimethylfluorescein (CDMFA)
5-sulfofluorescein diacetate (SFDA)
Cell Tracker Green – 5-chloromethylfluorescein diacetate (CTG)
5-(and-6)-carboxy-2´,7´-dichlorofluorescein diacetate, succinimidyl ester (DCF)
bis-(1,3-dibutylbarbituric acid)trimethine oxonol (DiBAC4(3))
3,3'-dipentyloxacarbocyanine iodide (DiOC5(3))
3,3'-dihexyloxacarbocyanine iodide (DiOC6(3))
Rhodamine 110

## Slide 65
All dyes
Oxonol Events: 3373
Human Lymphocytes
7 of 9 “green” dyes could be resolved

## Slide 66: Nanocrystals/Micro-Dots multiplexed systems
New probes
Potentially 1000’s of combinations
Sensitive, long lived, less bleaching
Difficult to make
Will require some advanced classification software

## Slide 67: Part 3
High Throughput Screening
and Advanced Analytical Systems

## Slide 69: Optical Design of a basic flow cytometer
PMT 1
PMT 2
PMT 5
PMT 4
Dichroic
Filters
Bandpass
Filters
Laser
Flow chamber
PMT 3
Scatter
Sensor
PMT 1
PMT 2
PMT 5
PMT 4
Dichroic
Filters
Bandpass
Filters
Laser
Flow chamber
PMT 3
Scatter
Sensor
Sample
HyperCyte diagram kindly supplied by Larry Sklar
HyperCytTM system added to flow cytometer
Copyright 2006 Nature Publishing Group
Nature Reviews Protocols
384 wells/10 min
1 ml/sample 5000 cells/ml
Notes: Basic schematic for how a flow cytometer works – a cell passes via a laser (blue line) and signals are passed down to each PMT and light scatter detector

## Slide 70
Automated
Cytomat
Incubator
Automated Biomek
Preparative system
Automated HyperCyte
Sample Delivery
Cyan flow cytometer

## Slide 71: Circa, 1990-present
Run 30-60 samples
per hour
ANALYSIS
1 sample 10 to 30 minutes

## Slide 73: The Operational Modality of Flow Cytometry
1. Take a complex biological sample
2. Separate that sample into single samples
3. Run each sample as a single unit
4. Put all the data back together again
5. Try to relate each piece of data back to the original biological system
The operational modality of Flow Cytometry.pptx
Traditional
HT Flow
1. Take a complex biological sample
2. Create very large scale “systems level” experimental grids
3. Run samples as a high speed, low volume set
4. Combine data automatically as sets
5. Conduct simultaneous
advanced analysis in
parallel

## Slide 74: Graphical data processing environments
Orange
IBM Data Explorer
Khoros
AVS

## Slide 75: Traditional gating process
Traditional flow gating logic and analysis approach
‘What we see depends mainly on what we are looking for.’ John Lubbock
Decision
Result
Decision
Decision
Result
Decision
Repeat ad infinitum

## Slide 76: Graphical cytometry analysis pipeline builder
A3
B3
C3
D3
E3
F3
G3
H3
I3
J3
Control +
Control -
Build curve
distance
Label 1
distance
Label 2
distance
Label 3
Extract features
We call this a Logic Map
Result
Drug 1
Drug 2
Drug 3
Drug 4
Drug 5
Drug6
Drug 7
Drug 8
Drug 9
Drug 10

## Slide 77: Rapid Processing
List Mode files
(96 or 384 wells)
(entire directory)
Protocol
IC50 or appropriate
result
About 30 seconds

## Slide 78: PlateAnalyzer Output
Different Models

## Slide 79: 3 functional assays, 32 drugs, 10 points per drug, 96 curves, (12 min to run 384 samples, 5000 cells per sample including all analysis
1 Click
1 minute
Mitosox
Calcein
mBBr

## Slide 81: Multiparametric and multifactorial HT cytometry
Define multiple populations in the feature space
Find the features describing the populations f1,.. fn
concentration
compounds
Create n × m array of features, where every m could be a vector [f1,…,fn]
HT flow cytometry
Find the features (such as IC50)
+
+

## Slide 82: Problem 1: eliminate operator’s input in analysis
gate
gate
Calculate fraction of “positive” and “negative” cells
Problems?
Well… plenty actually

## Slide 83: Alternative?
Method I
  Sort the apples (“gate” the green and red populations, or red and non-red)
  Count red and green (non-red)
  Compare the counts
Method II
  Find a bucket full of red apples only
  Find the distribution
  Take a bucket with red and green apples and find the distribution
  Compare the distribution
Harvest day
Percentage of green apples
-5
-10
-15
-20
20
40
60
80
Heuristic methods
  Minkowski-form
  Weighted-Mean-Variance (WMV)
Distance functions used in nonparametric tests
   2 (Chi Square)
  Kolmogorov-Smirnov (KS)
  Cramer/von Mises (CvM)
Information-theory divergences
  Kullback-Liebler (KL)
  Jeffrey divergence (JD)
Spectral measures
  Spectral angle mapper
  Bhatacharyya distance
Ground distance measures
  Histogram intersection (Overton)
  Quadratic form distance (QF)
  Wasserstein-Rubinstein-Mallows distance (Earth Movers Distance)

## Slide 84: Automated Gating
Purdue has developed current patent space:
  Distance measures for flow cytometry
  Variety of automated gating processes

## Slide 85: Analysis Logic Map Format

## Slide 86
Measuring KS Distance
Quadratic form Distance
Kullback-Leibler Distance
1. Select the desired parameter and press “DISTANCE”
2. A parameter box appears
3. The parameter box must be connected to both the drug box and the controls
You can select any parameter for KS distance, or quadratic form distance, or Kullback-Leibler distance if you wish. There are several different distance measures you can choose from – or add you own!!

## Slide 87: Comparison – Gating vs NonGating
IC50
KL-Values
*Patent pending

## Slide 88
CD3/CD14/CD7/cd38/Kappa/HLADR/CD57/CD11b/CD15/CD16/CD56/CD52/CD30/CD19/CD71
Different colors represent different phenotypes
Variable parameters
Variable parameters

## Slide 90: Mass Cytometry - CyTOF

## Slide 91
Data from Bernd Bodenmiller, Nolan Lab
dasatinib
H-89
imatinib
rapamycin
SB-202190
sorafenib
SP600125
staurosporine
sunitinib
UO126
VX680
BTK inhibitor III
IKK inhibitor X
JAK inhibitor I
JAK2 inhibitor III
JAK3 inhibitor VI
Lck inhibitor
PP2
Syk inhibitor IV
Akti-1/2,
GDC-0941
streptonigrin
Go-6983
Crassin
14x14x12=2353

## Slide 92: Very complex data sets

## Slide 93: PlateAnalyzer Logic Map for HC Analysis
(Example screenshot of the software)

## Slide 94: Cytospec Software PlateAnalyzer Software

## Slide 95: Cytospec Software
CYTOSPECTM (Written 2002 - updated Sept 16, 2012)
The following program is called Cytospec (TM) and is a general purpose flow cytometry analysis program
For 32 bit you need the EXE file (download now). For 64 bit use this EXE file (Download now).
Tutorial
Brief instructions for running Cytospec are here (download now) and will be updated one of these days!
1. Tutorial without audio (use above file) (play)
Video Tutorials: YOUTUBE & Twitter HANDLE is @CYTOMETRYMAN
The following video tutorials teach you all the basics of Cytospec. You can use the 5 color listmode file HERE if you want to play along!
1. Demo 1 -YOUTUBE or download- Basic introduction
2. Demo 2 - YOUTUBE or download-Tabloid output
3. Demo 3 - YOUTUBE or download- PCA - how to apply it\
4. Demo 4 - YOUTUBE or download- More PCA
5. Demo 5 - YOUTUBE or download- Some Utility Features
6. Demo 6 - YOUTUBE or download- More Utility Features
7. Demo 7 - YOUTUBE or download- Create a new FCS file and perform PCA on it

## Slide 96: Summary
Flow cytometry can be used for high throughput sampling for various drug screens, toxicology, etc
Flow cytometry can collect very high content data and new approaches to data processing are necessary
New approaches require more automation which leads to lower per test cost, better QC and faster results
New classification tools allow us to enhance the analytical process and assist in identifying signaling pathways
For analysis of these complex data sets – the complexity can be “under the hood”, but must be easy, interactive and accurate
Visualizing and manipulating very large data sets is quite possible using newly developed analytical tools